Hellenic Accreditation System

Annex F2/14 to the Certificate No. 473-5 SCOPE of ACCREDITATION of the Calibration Laboratory of ALEXIOU S.A.

| Calibration item / Parameters | Range of measurement | Expanded measurement uncertainty (k=2) * | Remarks |
| --- | --- | --- | --- |
| Mass measurements |
| Standard weights | 50 g | 2,0 mg | Calibration according to OΙΜL R111: 2004 Capability to calibrate standard weights up to (including) OIML class M2 |
|  | 100 g | 3,0 mg |  |
|  | 200 g | 4,0 mg |  |
|  | 500 g | 10 mg |  |
|  | 1 kg | 25 mg |  |
|  | 2 kg | 40 mg |  |
|  | 5 kg | 95 mg |  |
|  | 10 kg | 0,21 g |  |
|  | 20 kg | 0,37 g |  |
|  | 50 kg | 1,5 g |  |
|  | 100 kg | 2,0 g |  |
|  | 200 kg | 4,0 g |  |
|  | 500 kg | 10 g |  |
|  | 1000 kg | 20 g |  |
| Non-automatic weighing instruments (single – multiple scale), Weighbridges (continued) | 2 kg to 30000 kg | 1,6 x 10-4 | Calibration according to guide EURAMET /cg-18/v.04 (2015) Calibration is performed on-site. Available standards: OIML Class M2 in the range: 500 g to 1000 kg Max grouped load: 30000 kg |
|  | 30000 kg | 10 kg | Using substitution loads of known density, which are provided by the customer, according to instruction of the laboratory. |
|  | 40000 kg | 13 kg |  |
|  | 50000 kg to 60000 kg | 24 kg |  |

*	Where the expanded uncertainty (with coverage probability 95 %) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration & Measurement Capability (CMC) includes the measurand, the measurement range and the measurement uncertainty and expresses the smallest measurement uncertainty that can be achieved in a calibration.

Permanent laboratory premises address: 25th km Athens – Corinth New National Road, Mandra, GR-196 00 Attiki, Greece.

Approved Signatories: George Zafiris, Ioanna Bousoula, Maria Alexiou.

This Scope of Accreditation replaces the previous one dated 10.02.2021.

The Accreditation Certificate No. 473-5, to ELOT EN ISO/IEC 17025:2017, is valid until 13.11.2024.

Athens, 21.02.2024

Christos Nestoras

CEO of ESYD