Hellenic Accreditation System

Annex F2/22 to the Certificate No. 345-5 SCOPE of ACCREDITATION of the Calibration laboratory of algosystems S.A.

| Calibration item / Parameters | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Temperature measurements |  |  |  |
| Temperature block calibrators | [-80 oC to 0 oC] | 0,08 oC | According to guide EURAMET/cg-13/v.4:2017 |
|  | (0 oC to 270 oC] | 0,10 oC |  |
|  | (270 oC to 400 oC] | 0,20 oC |  |
|  | (400 oC to 1200 oC] | 1,40 oC |  |
| Thermocouples | [-40 oC to 140 oC] | 0,25 oC | According to guide EURAMET/cg-08/ v3.1:2020 |
|  | [140 oC to 400 oC] | 0,60 oC |  |
|  | [400 oC to 1200 oC] | 1,4 oC |  |
| Liquid-in-glass thermometers of total immersion | 35 oC to 90 oC | 0,16 oC | Internal Procedure ALGO-TEMP-0005, using water bath with a platinum resistance thermometer. |
| Climatic chambers, Furnaces, Ovens, Incubators, Autoclaves, Refrigerators, Freezers of volume up to 2000 L | -80 oC to 400 oC | 0,40 oC | According to guide DKD-R 5-7 09/2018, method C. Calibration can be performed on site. |
|  | [-25 oC to 70 oC] | 0,30 oC | According to guide DKD-R 5-7 09/2018 Calibration can be performed on site. |
|  | (70 oC to 400 oC] | 1,0 oC |  |
|  | (400 oC to 600 oC] | 1,4 oC |  |
| Chambers, Furnaces, Ovens | 400 oC to 1000 oC | 1,6 oC | Internal Procedure ALGO-TEMP-0003, using thermocouple type K, located at a point in the center of the chamber. Calibration can be performed on site. |
| Chambers, Storage rooms of volume > 2000 L | 10 oC to 70 oC | 0,40 oC | Internal Procedure ALGO-TEMP-0003, using platinum resistance thermometer at one point within the chamber. Calibration can be performed on site. |
|  |  | 1,0 oC | Internal Procedure ALGO-TEMP-0003, using type K thermocouples at one point within the chamber. Calibration can be performed on site. |
| Liquid baths | -80 oC to 300 oC | 0,10 oC | Internal Procedure ALGO-TEMP-0006, using platinum resistance thermometer. Calibration can be performed on site. |
| Dissolution testers, Disintegration testers | 35 °C to 45 °C | 0,10 oC | Internal Procedure ALGO-TEMP-0006, using platinum resistance thermometer at one point of sample placement. Calibration can be performed on site. |
| Analog – digital temperature recorders | [-40 oC έως 0 oC) | 0,18 oC | Internal Procedure ALGO-TEMP-0004, using climatic chamber and platinum resistance thermometer. |
|  | [0 oC έως 90 oC] | 0,15 oC |  |
|  | (90 oC έως 170 oC] | 0,46 oC |  |
| Temperature sensors with analog or digital display unit, Analog – digital temperature recorders | [-40 oC to 140 oC] | 0,10 oC | Internal Procedure ALGO-TEMP-0004, using temperature block calibrator and platinum resistance thermometer. Calibration can be performed on site. |
|  | (140 oC to 400 oC] | 0,50 oC |  |
|  | [50 oC to 200 oC] | 0,10 oC |  |
|  | [400 oC to 1200 oC | 1,7 oC | Internal Procedure ALGO-TEMP-0004, using temperature block calibrator and thermocouple type S. Calibration can be performed on site. |
| Temperature transmitter (span 100°C) with analog signal output (in mA or Volts), connected to temperature sensor | [-40 oC to 140 oC] | 0,10 oC | Internal Procedure ALGO-TEMP-00707, using temperature block calibrator and platinum resistance thermometer. Calibration can be performed on site. |
|  | (140 oC to 420 oC] | 0,50 oC |  |
|  | (420 oC to 1000 oC] | 1,7 oC | Internal Procedure ALGO-TEMP-0007, using temperature block calibrator and thermocouple type S. Calibration can be performed on site. |
| Optical Thermometers / Thermocameras | [-20 oC to 140 oC] | 0,30 oC | Internal Procedure ALGO-TEMP-0010 based on guidance VDI/VDE 3511 Part 4.4, using black body apparatus and platinum resistance thermometer. Calibration can be performed on site. |
|  | [140 oC to 250 oC] | 2,3 oC |  |
|  | (250 oC to 400 oC] | 3,5 oC |  |
| Thermal cyclers (PCR) | 50 oC to 100 oC | 0,12 oC | Internal Procedure ALGO-TEMP-0009 based on guidance ISO 20836:2021, using twenty temperature sensors and a temperature recording system. Calibration can be performed on site. |
| Humidity measurements |  |  |  |
| Analog – digital direct reading relative humidity measuring instruments humidity / Humidity recorders | 15 %RH to 95 %RH at 25 oC | 1,0 %RH to 1,4 %RH | Internal ProcedureALGO-HUM-0002, using climatic chamber and reference humidity sensor. |
| Climatic chambers of volume up to 2000 L | 15 %RH to 95 %RH at 25 oC | 1,1 %RH to 1,5 %RH | According to guide DKD-R 5-7 09/2018 Calibration can be performed on site. |
| Physical Measurements |  |  |  |
| Velocity / Fixed speed limit monitoring systems | 30 km/h to 250 km/h | 2,5 km/h | Internal Procedure ALGO-TRA-0001 Calibration can be performed on site. |
| Chemical Measurements |  |  |  |
| pH-meters | (0 pH to pH] | 0,022 pH | Internal Procedure ALGO-CHEM-0001 using standard reference solutions. Calibration can be performed on site. |
| Density Measurements |  |  |  |
| Hydrometers | 850 g/cm3 to 1200 g/cm3 | 6,0·10-4 g/cm3 to 6,5·10-4 g/cm3 | Internal Procedure ALGO-DENS-0001 based on Cuckow methodology. |
| Volume Measurements |  |  |  |
| Mechanical – digital Pipettes, Dispensers. | [0,1 μL to 1 μL] | 5,0 % | According to guide EURAMET/cg-19/ v3.0 (09/2018) and standards: ISO 8655-2:2022 ISO 8655-6:2022 |
|  | (1 μL to 2 μL] | 2,0 % |  |
|  | (2 μL to 10 μL] | 0,5 % |  |
|  | [10 μL to 100 μL] | 0,8 % to 0,21 % |  |
|  | (100 μL to 1000 μL] | 0,21 % to 0,12 % |  |
|  | (1000 μL to 10 mL] | 0,12 % |  |
| Titrators | [0,05 mL to 1 mL] | 0,9 % to 0,094 % |  |
|  | (1 mL to 10 mL] | 0,094 % to 0,024 % |  |
|  | (10 mL to 100 mL] | 0,024 % to 0,018 % |  |
| Siphons – Burettes, type «Ex» and «In» with volume V (ml), internal neck diameter D (mm) | [0,2 mL to 1 mL] / D=5mm | 1,1% to 0,1 % |  |
|  | [0,2 mL to 1 mL] / D=10mm | 4,5% to 0,5% |  |
|  | [1 mL to 10 mL] / D=5mm | 0,1% to 0,02 % |  |
|  | [1 mL to 10 mL] / D=10mm | 0,5% to 0,05 % |  |
|  | [1 mL to 10 mL] / D=15mm | 1,0% to 0,1 % |  |
|  | [10 mL to 50 mL] / D=5mm | 0,02% to 0,009 % |  |
|  | [10 mL to 50 mL] / D=10mm | 0,05% to 0,01 % |  |
|  | [10 mL to 50 mL] / D=15mm | 0,1% to 0,02 % |  |
|  | [10 mL to 50 mL] / D=20mm | 0,2% to 0,04 % |  |
|  | [10 mL to 50 mL] / D=30mm | 0,4% to 0,08 % |  |
|  | [50 mL to 100 mL] / D=5mm | 0,009% to 0,08 % |  |
|  | [50 mL to 100 mL] / D=10mm | 0,01% to 0,08 % |  |
|  | [50 mL to 100 mL] / D=15mm | 0,02% to 0,08 % |  |
|  | [50 mL to 100 mL] / D=20mm | 0,04% to 0,08 % |  |
|  | [50 mL to 100 mL] / D=30 mm | 0,08% to 0,08 % |  |
|  | [50 mL to 100 mL] / D=40 mm | 0,15% to 0,09 % |  |
| Glass - Plastic - Metallic Volumetric vessels, Flasks, Containers, type «Ex» and «In» with volume V (ml), and internal neck diameter D (mm) | [1 mL to 10 mL] / D=5 mm | 0,12 % to 0,016 % | Calibration according to standards and guides: ΑSΤΜ E542: 2022 ΟΙΜL R-120:2010 OIML R-43:1981 ISO 4787:2021 ISBN 978-963-88410-0-1:2009 IMEKO (p. 2344) |
|  | [1 mL to 10 mL] / D=10 mm | 0,47 % to 0,05 % |  |
|  | [1 mL to 10 mL] / D=15 mm | 1,1 % to 0,1 % |  |
|  | [1 mL to 10 mL] / D=20 mm | 1,9 % to 0,2 % |  |
|  | [10 mL to 100 mL] / D=5 mm | 0,016 % to 0,01 % |  |
|  | [10 mL to 100 mL] / D=10 mm | 0,05 % to 0,01 % |  |
|  | [10 mL to 100 mL] / D=15 mm | 0,1 % to 0,01 % |  |
|  | [10 mL to 100 mL] / D=20 mm | 0,2 % to 0,02 % |  |
|  | [10 mL to 100 mL]/ D=30 mm | 0,4 % to 0,04 % |  |
|  | [100 mLto 1 L] / D: (5 mm to 10 mm) | 0,01 % to 0,008 % |  |
|  | [100 mL to 1 L] / D=15 mm | 0,01 % to 0,008 % |  |
|  | [100 mL to 1 L]/ D=20 mm | 0,02 % to 0,009 % |  |
|  | [100 mL to 1 L] / D=30 mm | 0,04% to 0,009 % |  |
|  | [1 L to 5 L] / D: (5mm to 15mm) | 0,008 % to 0,037 % |  |
|  | [1 L to 5 L] / D:(20 mm to 30 mm) | 0,009 % to 0,037 % |  |
|  | [1 L to 5 L] / D:(40 mm to 60 mm) | 0,02 % to 0,037 % |  |
|  | [5 L to 20 L] / D=60 mm | 0,037 % to 0,032 % |  |
|  | [5 L to 20 L] / D=80 mm | 0,039 % to 0,032 % |  |
|  | [5 L to 20 L] / D=100 mm | 0,040 % to 0,032 % |  |
| Dimensional Measurements |  |  |  |
| Calipers, analog - digital | 0 mm to 300 mm d = (0,01 to 0,1) mm | 0,6 d | Internal Procedure ALGO-DIM-0001 Comparative calibration using gauge blocks and gauge rods. |
| Depth calipers | 0 mm to 300 mm d = (0,01 to 0,1) mm | 0,6 d |  |
| External Micrometers, analog - digital | 0 mm to 200 mm |  | Internal Procedure ALGO-DIM-0002 Comparative calibration using gauge blocks and gauge rods. |
|  | d = 0,001 mm | 0,6 μm to 3,7 μm |  |
|  | d = 0,01mm | 6 μm to 6,9 μm |  |
| Sieves | Nominal opening 20 μm to 125 mm | 3,5 μm with lens x50) 6,5 μm (with lens x30) | Calibration according to standards: ISO 3310-1:2016 ISO 3310-2:2013 ASTM E11:2022 ISO 5223:1995 EN 933-3:2012 |
| Profile projectors, Micrometers with scale | 0 mm to 50 mm | (d=0,0001mm) 0,87 μm (d=0,001mm) 1,1 μm (x30 < with lens< x50) | Internal Procedure ALGO-DIM-0004 Comparative calibration using glass microscales. Calibration can be performed on site. |
|  | 50 mm to 300 mm | (d=0,0001mm) 2,9 μm (d=0,001mm) 3,0 μm (with lens < x30) |  |
| Dial gauges | 0 mm to 25 mm |  |  |
|  | R = 2 μm | 3,6 μm | Internal Procedure ALGO-DIM-0006 Comparative measurement using dial indicator. |
|  | R = 5 μm | 4,0 μm |  |
|  | R = 10 μm | 6,4 μm |  |
|  | R = 20 μm | 12 μm |  |
|  | R = 100 μm | 58 μm |  |
|  | R = 5 μm | 13 μm | Internal Procedure ALGO-DIM-0006 Comparative measurement using dial indicator. Calibration can be performed on site. |
|  | R = 10 μm | 14 μm |  |
|  | R = 20 μm | 17 μm |  |
|  | R = 100 μm | 59 μm |  |
|  | 0 mm to 100 m | 0,5 μm | Internal Procedure ALGO-DIM-0006 Comparative measurement using gauge blocks. |
| Cubic moulds for concrete specimens 150x150x150 mm | Internal dimensions | 0,02 mm | Internal Procedure ALGO-DIM-0005 Tolerances according to EN 12390-1:2021 |
|  | Perpedicularity | 4 μm |  |
|  | Flatness | 3 μm |  |
| Cylindrical moulds (Standard proctor moulds) 4 in | Internal diameter: 101,6 mm or 152,4 mm Height A: 116,4 mm or 177,8 mm Height B: 50,8 mm | 0,02mm 0,03mm 0,04mm | Internal Procedure ALGO-DIM-0005 Tolerances according to ASTM D698:2021 §6.1.1 |
| Cylindrical moulds (Modified proctor moulds) 6 in |  |  | Internal Procedure ALGO-DIM-0005 Tolerances according to ASTM D698:2022, §6.1.2 |
| CBR cylindrical moulds (CBR moulds) 6 in |  |  | Internal Procedure ALGO-DIM-0005 Tolerances according to ASTM D1883:2021 §5.2 |
| Graduated rulers | 0 mm to 3000 mm | 0,25d d: minimum scale division in mm, 0,5mm | Internal Procedure ALGO-DIM-0005 based on OIML R35-2 : 2011 §7.5, §7.6 |
| Extensionmeters | 0 mm to 50 mm | 0,6μm + (0,85 μm per mm) | According to standards: ISO 9513:2012 ASTM E83:2016 Calibration can be performed on site. |
| Electrical measurements |  |  |  |
| DC Voltage (Generation) / DC Voltage Measuring Equipment | (0 mV to 329,9999 mV] | 1,8· 10-5 | Internal Procedure ALGO-DCLF-0001 based on the EURAMET cg-15 v3.0:2015, using standard multifunction calibrator. Calibration can be performed on site. |
|  | (0 V to 3,299999 V] | 8,3· 10-6 |  |
|  | (0 V to 32,99999 V] | 1,0· 10-5 |  |
|  | [30 V to 329,9999 V] | 1,5· 10-5 |  |
|  | [100 V to 1020,000 V] | 1,5· 10-5 |  |
| DC Voltage (Measurement) / DC Voltage Generators | [100 μV to 200 mV] | 1,5· 10-5 | Internal Procedure ALGO-DCLF-0006, using standard high-accuracy multimeter. Calibration can be performed on site. |
|  | (200 mV to 2 V] | 9,0· 10-6 |  |
|  | (2 V to 20 V] | 8,0· 10-6 |  |
|  | (20 V to 200 V] | 1,1· 10-5 |  |
|  | (200 V to 1000 V] | 1,3· 10-5 |  |
| DC Voltage (Measurement)/ DC Voltage Calibrators | 0,00001 mV to 100 mV] | 8,0· 10-6 | Internal Procedure ALGO-DCLF-0016, using standard multimeter (transfer standard) |
|  | (100mV to 1 V] | 4,0· 10-6 |  |
|  | (1V to 10 V] | 2,0· 10-6 |  |
|  | (10V to 100 V] | 3,0· 10-6 |  |
|  | (100V to 1000 V] | 4,0· 10-6 |  |
| AC Voltage (Generation) / AC Voltage Measuring Equipment | 1,0 mV to 32,999 mV [10 Hz to 45 Hz] | 8,1· 10-4 | Internal Procedure ALGO-DCLF-0002 using standard multifunction calibrator. Calibration can be performed on site. |
|  | (45 Hz to 10 kHz] | 2,8· 10-4 |  |
|  | (10 kHz to 20 kHz] | 2,8· 10-4 |  |
|  | (20 kHz to 50 kHz] | 9,3· 10-4 |  |
|  | (50 kHz to 100 kHz] | 9,4· 10-4 |  |
|  | (100 kHz to 500 KHz] | 3,0· 10-3 |  |
|  | 33,0 mV to 329,999 mV [10 Hz to 45 Hz] | 2,6· 10-4 |  |
|  | (45 Hz to 10 kHz] | 4,6· 10-4 |  |
|  | (10 kHz to 20 kHz] | 1,4· 10-4 |  |
|  | (20 kHz to 50 kHz] | 1,5· 10-4 |  |
|  | (50 kHz to 100 kHz] | 2,9· 10-4 |  |
|  | (100 kHz to 500 KHz] | 7,1· 10-4 |  |
|  | 0,33 V to 3,29999 V [10 Hz to 45 Hz] | 2,8· 10-4 |  |
|  | (45 Hz to 10 kHz] | 2,7· 10-4 |  |
|  | (10 kHz to 20 kHz] | 1,6· 10-4 |  |
|  | (20 kHz to 50 kHz] | 2,5· 10-4 |  |
|  | (50 kHz to 100 kHz] | 5,8· 10-4 |  |
|  | (100 kHz to 500 KHz] | 5,8· 10-4 |  |
|  | 3,3 V to 32,9999 V [10 Hz to 45 Hz] | 2,5· 10-4 |  |
|  | (45 Hz to 10 kHz] | 4,0· 10-4 |  |
|  | (10 kHz to 20 kHz] | 1,3· 10-4 |  |
|  | (20 kHz to 50 kHz] | 2,0· 10-4 |  |
|  | (50 kHz to 100 kHz] | 2,9· 10-4 |  |
|  | 33 V to 329,999 V 10 Hz to 45 Hz] | 2,1· 10-4 |  |
|  | (45 Hz to 10 kHz] | 1,5· 10-4 |  |
|  | (10 kHz to 20 kHz] | 1,7· 10-4 |  |
|  | (20 kHz to 50 kHz] | 2,7· 10-4 |  |
|  | (50 kHz to 100 kHz] | 2,7· 10-4 |  |
|  | 330 V to 1020 V 45 Hz to 1 kHz] | 2,6· 10-4 |  |
|  | (1 kHz to 5 kHz] | 2,0· 10-4 |  |
|  | (5 kHz to 10 kHz] | 2,1· 10-4 |  |
| AC Voltage (Generation) / AC Voltage Measuring Equipment | 10 mV to 200 mV [20 Hz to 40 Hz] | 1,3· 10-3 | Internal Procedure ALGO-DCLF-0007, using standard high-accuracy multimeter. Calibration can be performed on site. |
|  | (40 Hz to 1 kHz] | 5,0· 10-4 |  |
|  | (1 kHz to 10 kHz] | 6,2· 10-4 |  |
|  | (10 kHz to 30 kHz] | 3,6· 10-2 |  |
|  | (30 kHz to 100 kHz] | 3,6· 10-2 |  |
|  | (200 mV to 2 V] [20 Hz to 40 Hz] | 1,4· 10-3 |  |
|  | (40 Hz to 1 kHz] | 3,3· 10-4 |  |
|  | (1 kHz to 10 kHz] | 4,8· 10-4 |  |
|  | (10 kHz to 30 kHz] | 2,4· 10-3 |  |
|  | (30 kHz to 100 kHz] | 2,4· 10-3 |  |
|  | (100 kHz to 300 kHz] | 3,3· 10-2 |  |
|  | (300 kHz to 1 ΜΗz] | 3,6· 10-2 |  |
|  | (2 V to 20 V] [20 Hz to 40 Hz] | 1,3· 10-3 |  |
|  | (40 Hz to 1 kHz] | 3,3· 10-4 |  |
|  | (1 kHz to 10 kHz] | 3,6· 10-4 |  |
|  | (10 kHz to 30 kHz] | 2,4· 10-3 |  |
|  | (30 kHz to 100 kHz] | 2,4· 10-3 |  |
|  | (100 kHz to 300 kHz] | 3,6· 10-2 |  |
|  | (300 kHz to 1 ΜΗz] | 3,6· 10-2 |  |
|  | (20 V to 200 V [40 Hz to 1 kHz] | 3,3· 10-4 |  |
|  | (1 kHz to 10 kHz] | 3,6· 10-4 |  |
|  | (10 kHz to 30 kHz] | 2,4· 10-3 |  |
|  | (30 kHz to 100 kHz] | 2,5· 10-3 |  |
|  | (200 V to 700 V [20 Hz to 40 Hz] | 1,3· 10-3 |  |
|  | (40 Hz to 1 kHz] | 5,6· 10-4 |  |
| AC Voltage (Measurement)/ AC Voltage Calibrators | 1 mV [10 Hz to 1 kHz] | 4,2· 10-3 to 4,1· 10-3 | Internal Procedure ALGO-DCLF-0016, using standard multimeter (transfer standard). |
|  | [1 kHz to 100 kHz] | 4,1· 10-3 to 5,4· 10-3 |  |
|  | [100 kHz to 500 kHz] | 5,4· 10-3 to 2,2· 10-2 |  |
|  | [500 kHz to 1 ΜHz] | 2,2· 10-2 to 2,3· 10-2 |  |
|  | 10 mV [10 Hz to 1 kHz] | 6,0· 10-4 to 4,7· 10-4 |  |
|  | [1 kHz to 100 kHz] | 4,7· 10-4 to 9,0· 10-4 |  |
|  | [100 kHz to 500 kHz] | 9,0· 10-4 to 3,1· 10-3 |  |
|  | [500 kHz to 1 ΜHz] | 3,1· 10-3 to 4,4· 10-3 |  |
|  | 100 mV [10 Hz to 1 kHz] | 3,2· 10-4 to 1,4· 10-4 |  |
|  | [1 kHz to 100 kHz] | 1,4· 10-4 to 5,7· 10-4 |  |
|  | [100 kHz to 500 kHz] | 5,7· 10-4 to 1,4· 10-3 |  |
|  | [500 kHz to 1 ΜHz] | 1,4· 10-3 to 2,9· 10-3 |  |
|  | 1 V [10 Hz to 1 kHz] | 2,4· 10-4 to 4,5· 10-5 |  |
|  | [1 kHz to 100 kHz] | 4,5· 10-5 to 1,3· 10-4 |  |
|  | [100 kHz to 500 kHz] | 1,3· 10-4 to 1,0· 10-3 |  |
|  | [500 kHz to 1 ΜHz] | 1,0· 10-3 to 1,6· 10-3 |  |
|  | 10 V [10 Hz to 1 kHz] | 2,4· 10-4 to 4,2· 10-5 |  |
|  | [1 kHz to 100 kHz] | 4,2· 10-5 to 1,1· 10-4 |  |
|  | [100 kHz to 500 kHz] | 1,1· 10-4 to 1,0· 10-3 |  |
|  | [500 kHz to 1 ΜHz] | 1,0 ·10-3 to 1,6· 10-3 |  |
|  | 19 V, 1 kHz | 4,0· 10-5 |  |
|  | 100 V [10 Hz to 1 kHz] | 2,4· 10-4 to 5,1· 10-5 |  |
|  | [1 kHz to 100 kHz] | 5,1· 10-5 to 1,5· 10-4 |  |
|  | [100 kHz to 200 kHz] | 1,5· 10-4 to 8,6· 10-4 |  |
|  | 700V 50 kHz to 100 kHz | 2,5· 10-4 to 6,7· 10-4 |  |
|  | 1000 V [40 Hz to 1 kHz] | 7,9· 10-5 |  |
|  | [1 kHz to 30 kHz] | 7,9· 10-5 to 2,5· 10-4 |  |
| DC Current (Generation) / DC Current Measuring Equipment | (0 mA to 329.999 mA] | 2,0· 10-4 | Internal Procedure ALGO-DCLF-0003 based on EURAMET cg-15 v3.0:2015, using standard multifunction calibrator. Calibration can be performed on site. |
|  | (0 mA to 3.29999 mA] | 9,0· 10-5 |  |
|  | (0 mA to 32.9999 mA] | 9,2· 10-5 |  |
|  | (0 mA to 329.999 mA] | 9,2· 10-5 |  |
|  | (0 A to 1.09999 A] | 1,9· 10-4 |  |
|  | (1.1 A to 2.99999 A] | 3,1· 10-4 |  |
|  | (0 A to 10.0000 A] | 4,5· 10-4 |  |
| DC Current (Generation) / DC Current Clamp Meters | 2 Α | 5,1· 10-4 | Internal Procedure ALGO-DCLF-0012, using standard current source. Calibration can be performed on site. |
|  | 20 Α | 5,1· 10-4 |  |
|  | 100 Α | 7,1· 10-4 |  |
|  | 1000 Α | 6,4· 10-3 |  |
| DC Current (Measurement) / DC Current Calibrators up to 10 Α | (0,0001 μΑ to 100 μΑ] | 9,2· 10-5 | Internal Procedure ALGO-DCLF-0016, using standard multimeter (transfer standard). Calibration can be performed on site. |
|  | (100 μΑ to 1 mA] | 3,2· 10-5 |  |
|  | (1 mΑ to 10 mA] | 2,9· 10-5 |  |
|  | (10 mΑ to 100 mA] | 4,2· 10-5 |  |
|  | (100 mΑ to 1 A] | 6,2· 10-5 |  |
|  | (1 Α to 10 A] | 3,8· 10-4 |  |
| DC Current (Measurement) / DC Current Calibrators - Generators up to 100 Α | [2 mA to 20 mA] | 1,1· 10-4 | Internal Procedure ALGO-DCLF-0015, using standard current shunts and standard high-accuracy multimeter. Calibration can be performed on site. |
|  | [20 mA to 200 mA] | 1,1· 10-4 |  |
|  | [200 mA to 2 A] | 1,1· 10-4 to 2,1· 10-4 |  |
|  | [2 A to 20 A] | 2,1· 10-4 |  |
|  | [20 A to 100 A] | 2,1· 10-4 to 5,1· 10-4 |  |
| DC Current (Measurement) / DC Current Generators up to 2 Α | 100 μΑ to 2 A | 1,6· 10-4 | Internal Procedure ALGO-DCLF-0008, using standard high-accuracy multimeter. Calibration can be performed on site. |
| AC Current (Generation) / AC Current Measuring Equipment | 29,00 μA to 329,99 μA [10 Hz to 20 Hz] | 1,8· 10-3 | Internal Procedure ALGO-DCLF-0004 based on EURAMET cg-15 v3.0:2015, using a standard multifunction calibrator. Calibration can be performed on site. |
|  | (20 Hz to 45 Hz] | 1,6· 10-3 |  |
|  | (45 Hz to 1 kHz] | 1,4· 10-3 |  |
|  | (1 kHz to 5 kHz] | 2,7· 10-3 |  |
|  | (5 kHz to 10 kHz] | 7,0· 10-3 |  |
|  | (10 kHz to 30 kHz] | 9,2· 10-3 |  |
|  | 0,33 mA to 3,29999 mA [10 Hz to 20 Hz] | 1,6· 10-3 |  |
|  | (20 Hz to 45 Hz] | 1,0· 10-3 |  |
|  | (45 Hz to 1 kHz] | 8,1· 10-4 |  |
|  | (1 kHz to 5 kHz] | 2,1· 10-3 |  |
|  | (5 kHz to 10 kHz] | 6,7· 10-3 |  |
|  | (10 kHz to 30 kHz] | 8,4· 10-4 |  |
|  | 3,3 mA to 32,9999 mA [10 Hz to 20 Hz] | 1,6· 10-3 |  |
|  | (20 Hz to 45 Hz] | 1,0· 10-3 |  |
|  | (45 Hz to 1 kHz] | 8,1· 10-4 |  |
|  | (1 kHz to 5 kHz] | 1,6· 10-3 |  |
|  | (5 kHz to 10 kHz] | 3,9· 10-3 |  |
|  | (10 kHz to 30 kHz] | 7,9· 10-3 |  |
|  | 33 mA to 329,999 mA [10 Hz to 20 Hz] | 1,4· 10-3 |  |
|  | (20 Hz to 45 Hz] | 7,5· 10-3 |  |
|  | (45 Hz to 1 kHz] | 3,7· 10-4 |  |
|  | (1 kHz to 5 kHz] | 4,0· 10-4 |  |
|  | (5 kHz to 10 kHz] | 1,6· 10-3 |  |
|  | (10 kHz to 30 kHz] | 3,2· 10-3 |  |
|  | 0,33 A to 1,09999 A [10 Hz to 45 Hz] | 1,5· 10-3 |  |
|  | (45 Hz to 1 kHz] | 1,5· 10-3 |  |
|  | (1 t kHz o 5 kHz] | 4,8· 10-4 |  |
|  | (5 kHz to 10 kHz] | 2,0· 10-2 |  |
|  | 1,1 A to 2,09999 A [10 Hz to 45 Hz] | 1,5· 10-3 |  |
|  | (45 Hz to 1 kHz] | 5,0· 10-4 |  |
|  | (1 kHz to 5 kHz] | 5,0· 10-4 |  |
|  | (5 kHz to 10 kHz] | 2,1· 10-2 |  |
|  | 3 A to 10,9999 A [45 Hz to 100 Hz] | 6,3· 10-4 |  |
|  | (100 Hz to 1 kHz] | 1,3· 10-3 |  |
|  | (1 kHz to 5 kHz] | 1,6· 10-2 |  |
| AC Current (Generation) / AC Current Clamp Meters, 50 Hz | 2 Α | 2,5· 10-3 | Internal Procedure ALGO-DCLF-0012, using standard current generation device. Calibration can be performed on site. |
|  | 20 Α | 2,9· 10-3 |  |
|  | 100 Α | 4,7· 10-3 |  |
|  | 1000 Α | 6,4· 10-3 |  |
| AC Current (Measurement) / AC Current Calibrators up to 10 Α | 100 μΑ 10 Hz to 5 kHz | 3,7· 10-4 to 3,6· 10-4 | Internal Procedure ALGO-DCLF-0016, using standard multimeter (transfer standard) |
|  | 1 mA 10 Hz to 5 kHz | 2,5· 10-4 to 2,8· 10-4 |  |
|  | 10 mA 10 Hz to 5 kHz | 2,5· 10-4 to 2,3· 10-4 |  |
|  | 100 mA 10 Hz to 5 kHz | 2,5· 10-4 to 2,1· 10-4 |  |
|  | 1 A 20 Hz to 5 kHz | 2,7· 10-4 to 4,3· 10-4 |  |
|  | 10 A 40 Hz to 10 kHz | 3,9· 10-4 to 3,1· 10-3 |  |
| AC Current (Measurement) / AC Current Calibrators - Generators up to 100 Α | 2 mA to 20 mA 50 Hz to 1 kHz | 1,2· 10-3 | Internal Procedure ALGO-DCLF-0015, using standard current shunts and standard high-accuracy multimeter. Calibration can be performed on site. |
|  | 20 mA to 200 mA 50 Hz to 1 kHz | 1,2· 10-3 |  |
|  | 200 mA to 2 A 50 Hz to 1 kHz | 1,2· 10-3 |  |
|  | 2 A to 20 A 50 Hz to 1 kHz | 1,2· 10-3 |  |
|  | 20 A to 100 A 50 Hz to 1 kHz | 1,2· 10-3 |  |
| AC Current (Measurement) / AC Current Generators up to 2 Α | 10 mΑ to 2 A [30 Hz to 300 Hz] | 1,0· 10-3 | Internal Procedure ALGO-DCLF-0009, using standard high-accuracy multimeter. Calibration can be performed on site. |
|  | (300 Hz to 1 kHz] | 1,1· 10-3 |  |
|  | (1 kHz to 5 kHz] | 3,1· 10-3 |  |
| 4-wire Resistance (Generation) / Resistance Measuring Equipment | (0 Ω to 10,9999 Ω] | 5,8· 10-5 | Internal Procedure ALGO-DCLF-0005 based on EURAMET cg-15 v3.0:2015, using standard multifunction calibrator. Calibration can be performed on-site. |
|  | 11 to 32,9999 Ω | 2,4· 10-5 |  |
|  | 33 to 109,9999 Ω | 2,3· 10-5 |  |
|  | 110 Ω to 329,9999 Ω | 2,2· 10-5 |  |
|  | 330 Ω to 1,099999 kΩ | 2,2· 10-5 |  |
|  | 1,1 kΩ to 3,299999 kΩ | 2,2· 10-5 |  |
|  | 3,3 kΩ to 10,99999 kΩ | 2,2· 10-5 |  |
|  | 11 kΩ to 32,99999 kΩ | 2,2· 10-5 |  |
|  | 33 kΩ to 109,9999 kΩ | 2,2· 10-5 |  |
|  | 110 kΩ to 329,99999 kΩ | 2,7· 10-5 |  |
|  | 330 kΩ to 1,099999 MΩ | 2,6· 10-5 |  |
|  | 1,1 MΩ to 3,299999 MΩ | 4,7· 10-5 |  |
|  | 3,3 MΩ to 10,99999 MΩ | 4,8· 10-5 |  |
|  | 11 MΩ to 32,99999 MΩ | 2,0· 10-4 |  |
|  | 33 MΩ to 109,9999 MΩ | 4,0· 10-4 |  |
| 2-wire Resistance (Generation) / Resistance Measuring Equipment | 10 Ω | 1,3· 10-3 | Internal Procedure ALGO-DCLF-0005 based on the EURAMET cg-15 v3.0:2015, using standard multifunction calibrator. Calibration can be performed on-site. |
|  | 100 Ω | 2,0· 10-4 |  |
|  | 1 kΩ | 1,7· 10-4 |  |
|  | 10 kΩ | 2,4· 10-5 |  |
|  | 100 kΩ | 1,2· 10-5 |  |
|  | 1 ΜΩ | 2,4· 10-5 |  |
|  | 10 ΜΩ | 7· 10-4 |  |
|  | 100 ΜΩ | 8· 10-3 |  |
| Resistance / Insulation testers, Megohmmeters | 1 ΜΩ | 3,7 kΩ | Internal Procedure ALGO-DCLF-0019, using standard resistors. Calibration can be performed on-site. |
|  | 10 ΜΩ | 37 kΩ |  |
|  | 100 ΜΩ | 370 kΩ |  |
|  | 1 GΩ | 3,8 ΜΩ |  |
|  | 10 GΩ | 42 ΜΩ |  |
|  | 100 GΩ | 0,95 GΩ |  |
|  | 1 ΤΩ | 17 GΩ |  |
|  | 10 ΤΩ | 730 GΩ |  |
|  | 100 ΤΩ | 1,8 ΤΩ |  |
| Resistance (Measurement) / Resistance Calibrators | 10 Ω | 2,0· 10-5 | Internal Procedure ALGO-DCLF-0016, using standard multimeter (transfer standard). |
|  | 100 Ω | 8,0· 10-6 |  |
|  | 1 kΩ | 6,5· 10-6 |  |
|  | 10 kΩ | 6,5· 10-6 |  |
|  | 100 kΩ | 8,0· 10-6 |  |
|  | 1 ΜΩ | 1,3· 10-5 |  |
|  | 10 ΜΩ | 2,7· 10-5 |  |
|  | 100 ΜΩ | 8,5· 10-5 |  |
| 4-wire Resistance (Measurement) / Standard Resistors | [0,1 Ω to 200 Ω] | 1,6· 10-5 | Internal Procedure ALGO-DCLF-0010, using standard high-accuracy multimeter. Calibration can also be performed on-site. |
|  | (200 Ω to 2 kΩ] | 1,3· 10-5 |  |
|  | (2 kΩ to 20 kΩ] | 1,3· 10-5 |  |
|  | (20 kΩ to 200 kΩ] | 1,4· 10-5 |  |
|  | (200 kΩ to 2 ΜΩ] | 2,7· 10-5 |  |
|  | (2 ΜΩ to 20 ΜΩ] | 5,1· 10-5 |  |
|  | (20 ΜΩ to 200 ΜΩ] | 2,3· 10-4 |  |
| 2-wire Resistance (Measurement) / Standard Resistors | [0,1 Ω to 200 Ω] | 1,2· 10-3 | Internal Procedure ALGO-DCLF-0010, using standard high-accuracy multimeter. Calibration can also be performed on-site. |
|  | (200 Ω to 2 kΩ] | 2,1· 10-4 |  |
|  | (2 kΩ to 20 kΩ] | 2,8· 10-5 |  |
|  | (20 kΩ to 200 kΩ] | 1,6· 10-5 |  |
|  | (200 kΩ to 2 ΜΩ] | 2,8· 10-5 |  |
|  | (2 ΜΩ to 20 ΜΩ] | 5,1· 10-5 |  |
|  | (20 ΜΩ to 200 ΜΩ] | 2,3· 10-4 |  |
| Resistance (Measurement) / Current shunts | 0,001 Ω | 5,0· 10-4 | Internal ProcedureALGO-DCLF-0014, using standard current source and standard high-accuracy multimeter. Calibration can be performed on-site. |
|  | 0,01 Ω | 2,8· 10-4 |  |
|  | 0,1 Ω | 1,5· 10-4 |  |
|  | 1 Ω | 8,4· 10-5 |  |
|  | 10 Ω | 8,0· 10-5 |  |
|  | 100 Ω | 8,0· 10-5 |  |
| DC Voltage (Generation) / Temperature Indicators without cold junction compensation | ±100 mV -270 °C to 1700 °C | 0,07 oC | According to guide EURAMET/ cg-11/v.2:2011 Calibration can be performed on-site. |
| DC Voltage (Generation) / Temperature Indicators with cold junction compensation | ±100 mV -270 °C to 1700 °C | 0,27 oC |  |
| pH-meters (voltage measurement) | -410mV to 410mV | 0,01 mV | Internal Procedure ALGO-DCLF-0021 Calibration with voltage signal simulation. The measurement range corresponds to pH values ​​0 to 14. Calibration can be performed on-site. |
| Time / Frequency measurements |  |  |  |
| Timers | 0 s to 21.600 s | 0,23 s | Internal Procedure ALGO-DCLF-0018, using Rubidium frequency standard and standard frequency counter. |
| Frequency counters | 1 MHz | 14 μHz | Internal Procedure ALGO-DCLF-0017, using Rubidium frequency standard. |
|  | 10 MHz | 230 μHz |  |
| Optical tachometers | 600 rpm (10 Hz) | 14,2 · 10-2 rpm | Internal Procedure ALGO-DCLF-0013, using standard multifunction calibrator and standard frequency counter. Calibration can be performed on-site. |
|  | 900 rpm (15 Hz) | 14,3 · 10-2 rpm |  |
|  | 3.000 rpm (50 Hz) | 16 · 10-2 rpm |  |
|  | 15.000 rpm (250 Hz) | 23 · 10-2 rpm |  |
|  | 30.000 rpm (500 Hz) | 31 · 10-2 rpm |  |
|  | 60.000 rpm (1000 Hz) | 0,5 rpm |  |
|  | 90.000 rpm (1500 Hz) | 0,8 rpm |  |
| Centrifuges, Centrifuge extractors, Mixers | 0 rpm to 90 000 rpm | 3,0 rpm | Internal Procedure:ALGO-RPM—0001, using tachometer. Calibration can be performed on-site. |
| Mass Measurements |  |  |  |
| Standard weights |  |  | Capability to calibrate standard weights up to (including) OIML R 111-1:2004 Class: |
|  | [1 mg to 10 mg] | 0,010 mg | F2 |
|  | [20 mg to 100 mg] | 0,010 mg | F1 F1 |
|  | 200 mg | 0,010 mg |  |
|  | 500 mg | 0,013 mg |  |
|  | 1 g | 0,015 mg |  |
|  | 2 g | 0,020 mg |  |
|  | 5 g | 0,021 mg |  |
|  | 10 g | 0,025 mg |  |
|  | 20 g | 0,031 mg |  |
|  | 50 g | 0,040 mg |  |
|  | 100 g | 0,060 mg |  |
|  | 200 g | 0,12 mg 0,12 mg 0,12 mg |  |
|  | 500 g | 0,90 mg | F2 |
|  | 1 kg | 1,0 mg | F1 |
|  | 2 kg | 1,4 mg |  |
|  | 5 kg | 3,2 mg |  |
|  | 10 kg | 82 mg | Μ1 |
|  | 20 kg | 83 mg | F2 |
|  | 50 kg | 0,17 g |  |
| Non-automatic weighing instruments | 1 mg | 2 μg | According to guide EURAMET/cg-18/v.04:2015 Available Standard Weights OIML R 111-1:2004: Class E2 in the range from 1mg to 5kg. Max. grouped load: 11,1 kg Class F1 in the range from 1mg to 5kg. Max. grouped load: 11,1 kg Class F2 in the range from 1mg to 5kg. Max. grouped load: 11,1 kg Class F2 in the range from 10kg to 20kg. Max. grouped load: 50 kg Class M1 in the range from 10kg to 700kg. Max. grouped load: 700 kg Calibration is performed on-site. |
|  | 2 mg | 2 μg |  |
|  | 5 mg | 2 μg |  |
|  | 10 mg | 2 μg |  |
|  | 20 mg | 3 μg |  |
|  | 50 mg | 3 μg |  |
|  | 100 mg | 5 μg |  |
|  | 200 mg | 6 μg |  |
|  | 500 mg | 7 μg |  |
|  | 1 g | 10 μg |  |
|  | 2 g | 10 μg |  |
|  | 5 g | 15 μg |  |
|  | 10 g | 20 μg |  |
|  | 20 g | 30 μg |  |
|  | [50 g to 1 kg] | 1  10-6 |  |
|  | (12 kg to 5 kg] | 2  10-6 |  |
|  | (5 kg to 10 kg] | 6  10-6 |  |
|  | (10 kg to 60 kg] | 5  10-5 |  |
|  | (60 kg to 750 kg] | 6  10-5 |  |
|  | (700 kg to 6,84 t] | (5,3∙10-5m + 1,2∙N) kg Where: m: value of mass in kg N: number of substitution loads | According to guide EURAMET/cg-18/v.v.04:2015, using substitution loads of known density,which are provided by the client according to lab suggestion. Calibration is performed on-site. |
| Force measurements |  |  |  |
| Testing machines: |  |  | According to ISO 7500-1:2018 Calibration is performed on-site. |
| - Compression | (0,0001 kN to 1 kN] | 0,03 % | Class 0.5 machines |
|  | [2 kN to 20 kN] | 0,10% | ,, 1 ,, |
|  | [20 kN to 100 kN] | 0,05% | ,, 0.5 ,, |
|  | [100 kN to 450 kN] | 0,10% | ,, 1 ,, |
|  | [450 kN to 2000 kN] | 0,10 % | ,, 1 ,, |
|  | [2000 kN to 3000 kN | 0,35% | ,, 1 ,, |
| - Tension | (0,0001 kN to 1 kN] | 0,03% | Class 0.5 machines |
|  | [2 kN to 22 kN] | 0,08% | ,, 0.5 ,, |
|  | [20 kN to 100 kN] | 0,05% | ,, 0.5 ,, |
|  | [100 kN to 220 kN] | 0,05% | ,, 0.5 ,, |
|  | [220 kN to 450 kN] | 0,08% | ,, 0.5 ,, |
| Compression testing machines: |  |  | According to standards: ΕΝ 12390-4:2019 ISO 7500-1:2018 Calibration is performed on-site. |
| - Force transfer | Force ratio in load [200 kN to 2000 kN] | 0,01 |  |
| - Force rate | [1 kN/s to 100 kN/s] | 0,8 % |  |
| - Flatness of machine plates | [5 μm to 30 μm] | 3,5 μm |  |
| - Hardness of machine plates | [24 HRC to 63 HRC] | 1 HRC |  |
| - Roughness of machine plates | [0.2 Ra to 5 Ra] (μm) | 0.011 μm to 0.18 μm |  |
| Load rings, CBR, MARSHAL loading machines | (0,1 kN to 100 kN] | 0,23% to 0,05% | According to standards: ISO 7500-1: 2018 BS 1377-1:2016 Calibration is performed on-site. |
| Torque Measurements |  |  |  |
| Torque calibrators | [1 Nm to 85 Nm] | 0,25 % | Internal Procedure ALGO-TORQUE-0001, using a measuring axis and standard weights. |
|  | (85 Nm to 700 Nm] | 0,15 % |  |
| Torque meters | [1 Nm to 700 Nm] | 0,6 % | According to standards: ISO 6789-1:2017 ISO 6789-2:2017 using torque calibrators. |
| Pressure measurements |  |  |  |
| Relative pressure gauges of direct reading, analog – digital, Pressure recorders | [3,5 kPa to 100 kPa] | 3,3∙10-5∙P + 1,0∙10-4 (kPa) όπου P : τιμή πίεσης εκφρασμένη σε kPa | According to guide EURAMET/cg-17/v.4.1 (09/2022), using dead weight testers. Medium: gas. |
|  | [0,1 ΜPa to 10 MPa] | 1,4∙10-5∙P + 4,5∙10-4 (MPa) όπου P : τιμή πίεσης εκφρασμένη σε MPa |  |
|  | [0,5 MPa to 5,5 MPa] | 5,9∙10-5∙P + 8,1∙10-5 (MPa) όπου P : τιμή πίεσης εκφρασμένη σε MPa | According to guide EURAMET/cg-17/v.4.1 (09/2022), using dead weight testers. Medium: oil. |
|  | (5,5 MPa to 55 MPa] | 3,7∙10-5∙P + 9,5∙10-4 (MPa) Όπου P : τιμή πίεσης εκφρασμένη σε MPa |  |
|  | [-1000 Pa to 0 Pa] | 3,6 Pa | According to guide EURAMET/cg-17/v.4.1 (09/2022), using standard pressure transducers. Medium: gas. Calibration can be performed on-site. |
|  | [0 Pa to 1000 Pa] | 3,6 Pa |  |
|  | [10 kPa to 100 kpa] | 4,6 Pa | According to guide EURAMET/cg-17/v.4.1 (09/2022), using standard pressure transducers Medium: gas, oil. Calibration can be performed on-site. |
|  | [200 kPa to 2 Mpa] | 0,91 kPa to 1,1 kPa |  |
|  | [2 Mpa to 10,3 MPa] | 0,82 kPa |  |
|  | [10,3 MPa to 34,3 MPa] | 4,8 kPa to 5,4 kPa |  |
|  | [34,3 MPa to 100 MPa] | 7,9 kPa | According to guide EURAMET/cg-17/v.4.1 (09/2022), using standard pressure transducers. Medium: oil. Calibration can be performed on-site. |
| Relative negative pressure gauges of direct reading, analog – digital, Pressure recorders | [-95 kPa to 0 kPa] | 0,02 kPa | According to guide EURAMET/cg-17/v.4.1 (09/2022), using standard pressure transducers. Medium: gas. Calibration can be performed on-site. |
| Absolute pressure gauges of direct reading, analog – digital, Pressure recorders | [10 kPa to 100 kPa] | 4,6 Pa | According to guide EURAMET/cg-17/v.4.1 (09/2022), using standard pressure transducers. Medium: gas. Calibration can be performed on-site. |
|  | [200 kPa to 2 MPa] | 0,91 kPa to 1,1 kPa |  |
|  | [2 MPa to 10,3 MPa] | 0,82 kPa |  |
| Barometric Pressure | [900 hPa to 1050 hPa] | 0,25 hPa | According to guide EURAMET/cg-17/v.4.1 (09/2022), using standard pressure transducers. Medium: gas. Calibration can be performed on-site. |
|  |  |  | According to guide EURAMET/cg-17/v.4.1 (09/2022), using standard pressure transducers and pressure chamber. Medium: gas. |

* Where uncertainty (with a probability of 95% coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

Calibration & Measurement Capability (CMC), includes the measured quantity, measurement range and measurement uncertainty and expresses the smallest measurement uncertainty that can be achieved during a calibration.

Address of permanent laboratory facilities: 46, N. Plastira Str., GR-171 21, N. Smyrni, Attiki, Greece.

Approved Signatories: Τ. Manaras,  Μ. Tsoumpos, K. Mavroethakos,  M. Chalkia, K. Fotiou, G. Katsamakis, G. Kasapas, E. Athanasiadis.

This Scope of Accreditation replaces the previous one dated November 14th, 2023.

The Accreditation Certificate No. 345-4, to ELOT EN ISO/IEC 17025:2017, is valid until 21.05.2028.

Athens, February 5th, 2024

Christos Nestoras

CEO of ESYD