Hellenic Accreditation System

Annex F2/9 to the Certificate No. 966-4

SCOPE of ACCREDITATION

of the

Calibration Laboratory

Hellenic Armed Forces Calibration Agency (YPHD)

| Measurand / Calibration item | Range of measurement | Uncertainty * (Confidence level 95%) | Remarks |
| --- | --- | --- | --- |
| Dimensional Measurements |
| Calipers analog / digital (for external, internal and depth measurements) | 0 mm to 500 mm (0 in to 20 in) | L: length in mm d: resolution in μm | Comparative calibration using gauge blocks, standard bars and standard rings. Calibration according to: VDI/VDE/DGQ 2618 part 9.1: 2006 |
| External micrometers analog / digital | 0 mm to 500 mm (0 in to 20 in) | L: length in mm d: resolution in μm | Comparative calibration using gauge blocks and standard bars. Calibration according to: VDI/VDE/DGQ 2618 part 10.1: 2001 |
| Mass Measurements |
| Non-automatic weighing instruments, Electronic / Mechanical | 1 mg | 3,1 μg | According to Guide: EURAMET/cg-18/v.04 (2015) Using standard weights OIML Class E2 in range 1 mg ... 20 kg. Total max grouped load 300 kg. |
|  | 2 mg | 3,8 μg |  |
|  | 5 mg | 3,0 μg |  |
|  | 10 mg | 3,8 μg |  |
|  | 20 mg | 3,0 μg |  |
|  | 50 mg | 4,2 μg |  |
|  | 100 mg | 5,0 μg |  |
|  | 200 mg | 6,2 μg |  |
|  | 500 mg | 9,1 μg |  |
|  | 1 g | 10 μg |  |
|  | 2 g | 13 μg |  |
|  | 5 g | 19 μg |  |
|  | 10 g | 26 μg |  |
|  | 20 g | 28 μg |  |
|  | 50 g | 47 μg |  |
|  | 100 g | 77 μg |  |
|  | 200 g | 0,15 mg |  |
|  | 500 g | 0,34 mg |  |
|  | 1 kg | 0,63 mg |  |
|  | 2 kg | 1,2 mg |  |
|  | 5 kg | 2.9 mg |  |
|  | 10 kg | 5.6 mg |  |
|  | 20 kg | 11mg |  |
|  | 42 kg | 24 mg |  |
|  | 42 kg to 300 kg | 7,4x10-6 |  |
|  | 42 kg to 300 kg | (kg) Ι: load indication in kg d: resolution in kg n: number of substitute weights of reference mass mcl | Using substitute weights, each approximately equal to a reference mass mcl |
| Mass / Standard Weights | 1 mg | 2,0 μg | Standard Weights up to Class Ε2 according to OIML R-111:2004 |
|  | 2 mg | 2,0 μg | >> Ε2 >> |
|  | 5 mg | 2,0 μg | >> Ε2 >> |
|  | 10 mg | 2,5 μg | >> Ε2 >> |
|  | 20 mg | 3,0 μg | >> Ε2 >> |
|  | 50 mg | 4,0 μg | >> Ε2 >> |
|  | 100 mg | 5,0 μg | >> Ε2 >> |
|  | 200 mg | 6,0 μg | >> Ε2 >> |
|  | 500 mg | 8,0 μg | >> Ε2 >> |
|  | 1 g | 10 μg | >> Ε2 >> |
|  | 2 g | 12 μg | >> Ε2 >> |
|  | 5 g | 16 μg | >> Ε2 >> |
|  | 10 g | 20 μg | >> Ε2 >> |
|  | 20 g | 25 μg | >> E2 >> |
|  | 50 g | 30 μg | >> E2 >> |
|  | 100 g | 50 μg | >> E2 >> |
|  | 200 g | 0,10 mg | >> E2 >> |
|  | 500 g | 0,25 mg | >> E2 >> |
|  | 1 kg | 0,5 mg | >> E2 >> |
|  | 2 kg | 1,0 mg | >> E2 >> |
|  | 5 kg | 2,5 mg | >> E2 >> |
|  | 10 kg | 5,0 mg | >> E2 >> |
|  | 20 kg | 10 mg | >> E2 >> |
|  | 100g to 30kg | 5,0x10-7 | Standard Weights not classified according to OIML R-111:2004 |
| Temperature Measurements |
| Thermometers with analog or digital reading / Analog or digital temperature data loggers | -80 oC to -30 oC | 0,030 oC | Using bath and platinum resistance thermometer. Devices under calibration must be equipped with external sensor. Calibration according to: DKD-R 5-1:2018 |
|  | -30 oC to 150 oC | 0,022 oC |  |
|  | 150 oC to 250 oC | 0,028 oC |  |
| Electrical measurements |
| Resistance measuring instruments (Ohmmeters) (4 / 2 Wire) (Generate) | 0 Ω | 50 μΩ | Using multifunction calibrator. According to guide: Euramet/cg-15/v.3.0.(02/2015) |
|  | 1 Ω | 11×10-5 |  |
|  | 1,9 Ω | 11×10-5 |  |
|  | 10 Ω | 29×10-6 |  |
|  | 19 Ω | 29×10-6 |  |
|  | 100 Ω | 14×10-6 |  |
|  | 190 Ω | 14×10-6 |  |
|  | 1 kΩ | 11×10-6 |  |
|  | 1,9 kΩ | 11×10-6 |  |
|  | 10 kΩ | 11×10-6 |  |
|  | 19 kΩ | 11×10-6 |  |
|  | 100 kΩ | 15×10-6 |  |
|  | 190 kΩ | 15×10-6 |  |
|  | 1 MΩ | 27×10-6 |  |
|  | 1,9 MΩ | 28×10-6 |  |
|  | 10 MΩ | 52×10-6 |  |
|  | 19 MΩ | 60×10-6 |  |
|  | 100 MΩ | 24×10-5 |  |
|  | 1 mΩ | 36×10-6 | Using current shunt. According to guide: Euramet/cg-15/v.3.0.(02/2015) |
|  | 10 mΩ | 10×10-6 |  |
|  | 100 mΩ | 3,9×10-6 |  |
|  | 1 Ω | 1,1×10-6 |  |
|  | 10 Ω | 1,8×10-6 |  |
|  | 100 Ω | 0,7×10-6 |  |
|  | 1 kΩ | 5,9×10-6 |  |
|  | 10 kΩ | 7,7×10-6 |  |
|  | 0,01 Ω | 24×10-5 | Using decade resistance. According to guide: Euramet/cg-15/v.3.0.(02/2015) |
|  | 0,1 Ω | 64×10-6 |  |
|  | 1 Ω | 20×10-6 |  |
|  | 10 Ω | 3,9×10-6 |  |
|  | 100 Ω | 1,0×10-6 |  |
|  | 1 kΩ | 3,5×10-6 |  |
|  | 10 kΩ | 4,2×10-6 |  |
|  | 100 kΩ | 16×10-6 |  |
| Standard Resistors & Current Shunts (4 / 2 wire) (Measure) |  |  | Using resistance measurement system. Max applied current 2 A. |
|  | 1 mΩ | 114×10-6 |  |
|  | 10 mΩ | 11×10-6 |  |
|  | 100 mΩ | 3,6×10-6 |  |
|  | 1 Ω | 0,8×10-6 |  |
|  | 10 Ω | 1,1×10-6 |  |
|  | 100 Ω | 0,8×10-6 |  |
|  | 1 kΩ | 1,4×10-6 |  |
|  | 10 kΩ | 1,4×10-6 |  |
|  | 100 kΩ | 4,8×10-6 |  |
|  | 1 MΩ | 9,6×10-6 |  |
|  | 10 MΩ | 15×10-6 |  |
|  | 100 MΩ | 15×10-6 |  |
| Resistance (2-wire) (Measure) | 10 Ω | 25×10-3 | Using digital multimeter 8 ½ digit. |
|  | 100 Ω | 2,5×10-3 |  |
|  | 1 kΩ | 0,26×10-3 |  |
|  | 10 kΩ | 36×10-6 |  |
|  | 100 kΩ | 13×10-6 |  |
|  | 1 MΩ | 17×10-6 |  |
|  | 10 MΩ | 60×10-6 |  |
|  | 100 MΩ | 51×10-5 |  |
|  | 1GΩ | 5,0×10-3 |  |

* Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative

Site of assessment: Permanent laboratory premises, Hellenic Air Force Police Military Camp, 16233, Vyronas

Approved Signatories: Ε. Artsitas, N. Andrianakis, T. Anagnostopoulos

This scope of Accreditation replaces the previous one dated 15.02.2023.

The Accreditation Certificate No. 966-4, to ELOT EN ISO/IEC 17025:2017, is valid until 22.01.2028.

Athens, 06.03.2023

Christos Nestoras

CEO of ESYD