Hellenic Accreditation System

Annex F2/13 to the Certificate No. 559-5

SCOPE of ACCREDITATION of the Calibration Laboratory of DEKA S.A.

| Measurand / Calibration item | Range of measurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass Μeasurements |
| Mass / Non-automatic weighing instruments | 1 mg | 5,3 μg | According to guideline: EURAMET cg-18/v.04 (11/2015) Using standard weights according to OIML κατά OIML Ρ-111-1 (2004): : E2: 1 mg … 200 g Max grouped load 611,11 g F1: 500 g … 5 kg Max grouped load 11,0 kg F2: 10 kg … 20 kg Max grouped load 30,0 kg M1: 20 kg Max grouped load 280 kg Calibration is performed on-site. |
|  | 2 mg | 5,3 μg |  |
|  | 5 mg | 5,3 μg |  |
|  | 10 mg | 6,9 μg |  |
|  | 20 mg | 8,1 μg |  |
|  | 50 mg | 11 μg |  |
|  | 100 mg | 13 μg |  |
|  | 200 mg | 16 μg |  |
|  | 500 mg | 21 μg |  |
|  | 1 g | 27 μg |  |
|  | 2 g | 35 μg |  |
|  | 5 g | 43 μg |  |
|  | 10 g | 53 μg |  |
|  | 20 g | 70 μg |  |
|  | 50 g | 81 μg |  |
|  | 100 g | 0,13 mg |  |
|  | 200 g | 0,27 mg |  |
|  | 500 g | 1,1 mg |  |
|  | 1 kg | 4,4 mg |  |
|  | 2 kg | 11 mg |  |
|  | 5 kg | 23 mg |  |
|  | 10 kg | 43 mg |  |
|  | 20 kg | 0,28 g |  |
|  | 50 kg | 1.3 g |  |
|  | 100 kg | 3.8 g |  |
|  | 200 kg | 8.2 g |  |
|  | 300 kg | 13 g |  |
| Temperature measurements |
| Temperature/ Analog and digital direct reading thermometers | [-78,5] oC | 0,13 oC | Calibration method based on COFRAC LAB GTA 08: 2012. |
|  | [-40 ... 250) oC | 0,047 oC |  |
|  | [250 ... 420) oC | 0,17 oC |  |
|  | [420 ... 600) oC | 1,3 oC |  |
|  | [600 ... 1000] oC | 1,8 oC |  |
|  | [-78,5] oC | 0,13 oC | Calibration method based on COFRAC LAB GTA 08: 2012. Calibration is performed on site. |
|  | [-30 ... 150) oC | 0,10 oC |  |
|  | [150 ... 420] oC | 0,17 oC |  |
| Temperature/ Liquid-in-glass thermometers of total and partial immersion | [-40 ... 250] oC | 0,047 oC | Calibration method based on NIST guide “Liquid in – Glass Thermometer Calibration Service”, 1988 |
| Temperature/ Analog and digital direct reading thermometers for measurements of air temperature, temperature data loggers | [-40 ... 5) oC | 0,36 oC | Comparison method. |
|  | [5 ... 15) oC | 0,31 oC | Comparison method. Calibration can also be performed on site. |
|  | [15 ... 35) oC | 0,16 oC |  |
|  | [35 ... 50] oC | 0,27 oC |  |
|  | (50 ... 70] oC | 0,36 oC | Comparison method. |
| Temperature / Temperature transmitters with analog signal output in mA or V connected with temperature sensor | [-40 oC ... 250] oC | 0,050 oC | Calibration method based on DKD-R 5-1:2018. |
| Temperature / Temperature transmitters with analog signal output in mA or V connected with temperature sensor for measurements of air temperature | [-40 ... 5) oC | 0,40 oC | Calibration method based on DKD-R 5-1:2018. |
|  | [5 ... 15) oC | 0,31 oC |  |
|  | [15 ... 35) oC | 0,16 oC |  |
|  | [35 ... 50] oC | 0,27 oC |  |
|  | (50 ... 70] oC | 0,40 oC |  |
| Temperature/ Climatic and isothermal chambers, Ovens Incubators and sterilization furnaces Refrigerators, Freezers, with internal volume up to 2000 l, with or without air circulation | [-40 ... 20] oC | 0,20 oC | According to DKD-R 5-7: 2018 using Pt 100 sensors. Calibration is performed in the whole volume or at selected points of the chamber. Calibration can also be performed on site. |
|  | [20 ... 140] oC | 0,10 oC |  |
|  | [140 ... 250] oC | 0,20 oC |  |
| Relative humidity measurements |
| Relative humidity / Analog and digital hygrometers, humidity data loggers Humidity transmitters with analog signal output in mA or V, connected with humidity sensor | [15 … 30) % RH | 1,0 % RH | Calibration method based on NPL guide “A guide to the measurement of humidity”, 1996. Calibration is performed at 25 oC. Calibration can also be performed on site. |
|  | [30 … 75) % RH | 1,3 % RH |  |
|  | [75 … 95] % RH | 1,6 % RH |  |
| Relative humidity / Analog and digital hygrometers, humidity data loggers Humidity transmitters with analog signal output in mA or V, connected with humidity sensor | [15 … 30) % RH | 1,0 % RH | Calibration method based on NPL guide “A guide to the measurement of humidity”, 1996. Calibration is performed at 40 oC. Calibration can also be performed on site. |
|  | [30 … 75) % RH | 1,5 % RH |  |
|  | [75 … 95] % RH | 1,8 % RH |  |
| Pressure measurements |
| Pressure/ Analog and digital positive or negative relative pressure gauges of direct reading (gas medium) | [-1000 … -125] Pa | 1,2 Pa | Calibration according to DKD-R 6-1:2014. Calibration can also be performed on site. |
|  | [-100 … -12,5] Pa | 0,35 Pa |  |
|  | (12,5 …100] Pa | 0,35 Pa |  |
|  | [125…1000] Pa | 1,2 Pa |  |
|  | [3 ... 25) kPa | 1,3 Pa |  |
|  | [25 ... 40) kPa | 20 Pa |  |
|  | [40 ... 130) kPa | 11 Pa |  |
|  | (130 … 400) kPa | (8·10-5·p) kPa όπου p: πίεση σε kPa |  |
|  | [400 … 1250) kPa | 0,10 kPa |  |
|  | (1,25 … 4,0] MPa | (8·10-2·p) kPa όπου p: πίεση σε MPa |  |
| Pressure/ Analog and digital relative pressure gauges of direct reading (deionized water medium) | [1,25 … 4) MPa | (8·10-2·p) kPa όπου p: πίεση σε MPa | Calibration according to DKD-R 6-1:2014. Calibration can also be performed on site. |
|  | [4 … 6) MPa | 3,0 kPa |  |
|  | [6 … 10) MPa | 5,0 kPa |  |
|  | [ 10 … 16) MPa | 8,0 kPa |  |
|  | [16 … 20) MPa | 10 kPa |  |
|  | [20… 40) MPa | 20 kPa |  |
|  | [40 … 100] MPa | 50 kPa |  |
| Pressure/ Analog and digital relative pressure gauges of direct reading (oil medium) | [1 … 10) MPa | (6,9·10-5·p + 1,4·10-4) MPa όπου p: πίεση σε MPa | Calibration according to DKD-R 6-1:2014, using pressure balance (DWT). |
|  | [10 … 100] MPa | (1,3·10-4·p – 1,0·10-3) MPa όπου p: πίεση σε MPa |  |
| Pressure/ Analog and digital absolute pressure gauges, including barometers, of direct reading (gas medium) | [5 ... 80) kPa abs | 80 Pa | Calibration according to DKD-R 6-1:2014. |
|  | [80 … 120) kPa abs | 20 Pa |  |
|  | [120 … 160) kPa abs | 80 Pa |  |
|  | [0,16 … 1,6] MPa abs | 0,80 kPa |  |

* Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

Site of assessment: Permanent laboratory premises, 40 Polydefkous str., 18545 Piraeus, Greece.

Approved Signatories: S. Kritikos, A. Limniatis, P. Mermigkas, V. Paroutsas

This Scope of Accreditation replaces the previous one dated 31.05.2021.

The Accreditation Certificate No. 559-5, to ELOT EN ISO/IEC 17025:2017, is valid until 31.05.2025.

Athens, 30.09.2021

Spyridon Podaras

CEO of ESYD