Hellenic Accreditation System

Annex F2/10 to the Certificate No. 941-6

SCOPE of ACCREDITATION of the Calibration Laboratory of MAIN SOLUTIONS Ltd for the performance of calibrations

| Measurand / Calibration item | Range of measurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass Μeasurements |
| Mass / Standard weights | 5 kg | 48 mg | According to guideline OIML R111-1: 2004. Calibration capability of standard weights up to class M1 acc. to OIML R111. |
|  | 10 kg | 55 mg |  |
|  | 20 kg | 0,11 g |  |
| Mass / Non-automatic weighing instruments | 1 mg | 1,5 μg | According to guideline EURAMET cg-18/v.04 (2015) Using standard weights of OIML R111 class: Ε1: 1 mg to 200g Max grouped load 611 g E2: 1mg to 20 kg Max grouped load 41722 g F1: 500 g to 5 kg Max grouped load 10,5 kg M1: 1 kg to 20 kg Max grouped load 288 kg Calibration is performed on-site. |
|  | 2 mg | 1,5 μg |  |
|  | 5 mg | 1,5 μg |  |
|  | 10 mg | 1,7 μg |  |
|  | 20 mg | 1,9 μg |  |
|  | 50 mg | 2,3 μg |  |
|  | 100 mg | 2,8 μg |  |
|  | 200 mg | 3,6 μg |  |
|  | 500 mg | 4,5 μg |  |
|  | 1 g | 5,5 μg |  |
|  | 2 g | 7,1 μg |  |
|  | 5 g | 9,0 μg |  |
|  | 10 g | 14 μg |  |
|  | 20 g | 16 μg |  |
|  | 50 g | 19 μg |  |
|  | 100 g | 28 μg |  |
|  | 200 g | 55 μg |  |
|  | 500 g | 0,90 mg |  |
|  | 1 kg | 2,6 mg |  |
|  | 2 kg | 5,1 mg |  |
|  | 5 kg | 0,12 g |  |
|  | 10 kg | 0,25 g |  |
|  | 20 kg | 0,46 g |  |
|  | 100 kg | 2,3 g |  |
|  | 150 kg | 4,1 g |  |
|  | 200 kg | 6,5 g |  |
|  | 300 kg | 11 g |  |
| Temperature measurements |
| Temperature / Temperature controlled chambers, ovens, incubators, autoclaves, refrigerators, freezers with volume up to 2000l, with or without air circulation | [-80 °C to -30 (C) | 0,30 °C | EURAMET cg 20/v.5.0 (2017) DKD-R 5-7: 09/2018 Calibration can also be performed on-site. |
|  | [-30 °C to +140 (C] | 0,095 °C |  |
|  | (140 °C to 250 (C] | 0,32 °C |  |
|  | (250 °C to 350 (C] | 0,47 °C |  |
| Temperature / Ovens, furnaces with volume up to 2000l, with or without air circulation | 200 (C to 500( C | 0,5 °C to 1,3 °C | EURAMET cg 20/v.5.0 (2017) DKD-R 5-7: 09/2018, case (C) Calibration can also be performed on-site. |
| Temperature / Liquid baths | [-80 °C to -30 (C) | 0,30 °C | EURAMET cg 20/v.5.0 (2017) DKD-R 5-7: 09/2018 Calibration can also be performed on-site. |
|  | [-30 °C to +140 (C] | 0,095 °C |  |
|  | (140 °C to 250 (C] | 0,32 °C |  |
|  | (250 °C to 350 (C] | 0,47 °C |  |
| Temperature / Dissolution testers and disintegration testers | 35 °C to 40 °C | 0,27 °C | Internal method G.05.04 using Pt100 sensors and K-type thermocouples. Calibration can also be performed on-site. |
| Temperature / Thermal Cyclers (PCR), Block calibrators, Thermal reactors | [-30 °C to 0 (C) | 0,44 °C | Internal method G.05.04 using Pt100 sensors and K-type thermocouples. Calibration can also be performed on-site. |
|  | [0 °C to 70 (C) | 0,27 °C |  |
|  | [70 °C to 100 (C) | 0,36 °C |  |
|  | [100 °C to 150 (C) | 0,53 °C |  |
|  | [150 °C to 200 (C] | 0,92 °C |  |
| Temperature / Digital or analogue direct reading thermometers | [-80 °C to 140 °C] | 0,07 °C | Internal method G.05.01 using dry block calibrator and Pt100 sensor. Calibration can also be performed on-site. |
|  | (140 °C to 200 °C] | 0,15 °C |  |
|  | (200 °C to 350 °C] | 0,20 °C |  |
| Temperature / Temperature recorders, data loggers | [-20 °C to 0 °C) | 0,37 °C | Internal method G.05.01 using Pt100 sensor and climatic chamber. |
|  | [0 °C to 70 (C) | 0,28 °C |  |
|  | [70 °C to 120 (C) | 0,30 °C |  |
|  | [120 °C to 150 (C) | 0,36 °C |  |
|  | [150 °C to 200 °C) | 0,45 °C |  |
| Relative Humidity measurements |
| Relative humidity / Analog and digital hygrometers | [10 %rh to 50 %rh) [50 %rh to 70 %rh) [70 %rh to 90 %rh] at 25 °C | 1,2 %rh 1,4 %rh 1,7 %rh | Internal method G.05.03 using reference hygrometer in climatic chamber and temperature/humidity calibrator. Calibration can also be performed on-site. |
|  | [20 %rh to 30 %rh) [30 %rh to 50 %rh) [50 %rh to 75 %rh] at 40 °C | 1,1 %rh 1,2 %rh 1,3 %rh |  |
| Relative humidity / Climatic chambers with volume up to 2000l | 20 %rh to 90 %rh at 25 οC | 2,0 %rh | EURAMET cg 20/v.5.0 (2017) DKD-R 5-7: 09/2018 Calibration can also be performed on-site. |
|  | 30 %rh to 90 %rh at 40 οC | 2,3 %rh |  |
| Volume Μeasurements |
| Volume / Piston type pipettes | [2 μL to 20 μL] | 4,0 % to 2,5 % | EURAMET cg 19/v.3.0 (2018) |
|  | [20 μL to 100 μL] | 2,5 % to 0,3 % |  |
|  | [100 μL to 1000 μL] | 0,3 % to 0,16 % |  |
|  | [1000 μL to 10 mL] | 0,16 % |  |
| Volume / Piston type burettes (dispensers, digital burettes, automatic titrators) | 1 mL to 50 mL | 0,16 % of nominal volume |  |
| Volume / Burettes | [0,5 mL to 10 mL] | 0,13 % of maximum volume | EURAMET cg 19/v.3.0 (2018) |
|  | (10 mL to 100 mL] | 0,1 % of maximum volume |  |
| Volume / Volumetric flasks | 5 mL | 0,04 % | EURAMET cg 19/v.3.0 (2018) |
|  | 10 mL | 0,05 % |  |
|  | 20 mL | 0,05 % |  |
|  | 25 mL | 0,05 % |  |
|  | 50 mL | 0,05 % |  |
|  | 100 mL | 0,05 % |  |
|  | 200 mL | 0,05 % |  |
|  | 250 mL | 0,05 % |  |
|  | 500 mL | 0,05 % |  |
| Volume / Volumetric Cylinders | [0,5 mL to 50 mL] | 0,6 % of nominal volume | EURAMET cg 19/v.3.0 (2018) |
|  | (50 mL to 100 mL] | 0,3 % of nominal volume |  |
|  | (100 mL to 250 mL] | 0,1 % of nominal volume |  |
|  | (250 mL to 500 mL] | 0,2 % of nominal volume |  |
| Volume / Pipettes (fixed volume) | 0,5 mL to 100 mL | 0,05 % of nominal volume | EURAMET cg 19/v.3.0 (2018) |
| Volume / Pipettes (graduated) | [0,1 mL to 0,5 mL] | 1,2 % of nominal volume | EURAMET cg 19/v.3.0 (2018) |
|  | (0,5 mL to 20 mL] | 0,6 % of nominal volume |  |
|  | (20 mL to 100 mL] | 0,3 % of nominal volume |  |
| Pressure Measurements |
| Pressure / Analog and digital gauge pressure measuring instruments of direct reading | -94 kPa to 0 kPa | 0,27 kPa | DKD-R 6-1 03/2014 Gas pressure medium. |
|  | 49 kPa to 4 MPa | 0,4 kPa |  |
|  | [49 kPa to 4 MPa) | 0,4 kPa | DKD-R 6-1 03/2014 Oil or water pressure medium. |
|  | [4 MPa to 7 MPa) | 0,8 kPa |  |
|  | [7 MPa to 68 MPa] | 24 kPa |  |
| Torque Measurements |
| Torque meters / wrenches | 7,5 Nm to 150 Nm | 0,8 % | According to ISO 6789-1:2017 ISO 6789-2:2017 using torque calibrator. Only clockwise measurements (CW). |
|  | 150 Nm to 1500 Nm | 0,4 % |  |
| Spectrophotometric Measurements |
| Absorption / Visible and ultraviolet spectrophotometers in the range: 230 nm to 650 nm | Α ≤ 0,5 AU | 5,0 mAU ** | ASTM E 275: 2008 (Re 2022) (baseline stability, photometric accuracy and precision) Calibration is also performed on-site. |
|  | 0,5 < Α ≤ 1,0 AU | 6,2 mAU ** |  |
| Time / Frequency measurements |
| Time / Timers, stopwatches | 0 s to 20 h | 0,3 s | According to NIST SP 960-12 Stopwatch and Timer Calibrations (2009 edition) using calibrated timer. |
| Frequency / Non-contact optical tachometers | [6 to 1000] rpm | 0,06 % | Internal method G.03.02 using reference tachometer calibrator. |
|  | (1000 to 10000] rpm | 0,007 % |  |
|  | (10000 to 90000] rpm | 0,002 % |  |
| Frequency / Contact tachometers | [6 to 100] rpm | 0,5 % | Internal method G.03.02 using tachometer calibrator and reference frequency meter. |
|  | (100 to 1000] rpm | 0,1 % |  |
|  | (1000 to 10000] rpm | 0,02 % |  |
| Rotation frequency / Centrifuges, Agitators, Mixers, Homogenizers, Motors | [6 to 1000] rpm | 0,1 % | Internal method G.03.02 using reference frequency meter. Calibration can also be performed on-site. |
|  | (1000 to 90000] rpm | 0,01 % |  |

* Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

** Uncertainty values are estimated assuming a spectrophotometer having a resolution of 0.1mAU.

Site of assessment: Permanent laboratory premises, 38 Argostoliou str., 18 345 Moschato, Greece.

Approved Signatories: Theodoros Chilas, Serafeim Golias, Vasilios Papantinas, Nikolaos Dalakouras, Alexios Kontogeorgis, Stefanos Ramfos, Dimitrios Tampakopoulos.

This Scope of Accreditation replaces the previous one dated 18.10.2022.

The Accreditation Certificate No. 941-6, to ELOT EN ISO/IEC 17025: 2017, it is valid until 04.11.2027.

Athens, May 2nd, 2023

Christos Nestoras

CEO of ESYD