Hellenic Accreditation System

Annex F2/10 to the Certificate No. 900-5 SCOPE of ACCREDITATION Of the Calibration Laboratory METROCAL PC

| Calibration item / Parameters | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass Measurements |
| Standard weights |  |  | According to guide: OIML R111-1:2004 Calibration capability of standard weights up to (including) OIML class: |
|  | 1 mg | 2,0 μg | Ε2 |
|  | 2 mg | 2,0 μg | Ε2 |
|  | 5 mg | 2,0 μg | Ε2 |
|  | 10 mg | 2,5 μg | Ε2 |
|  | 20 mg | 3,3 μg | Ε2 |
|  | 50 mg | 4,0 μg | Ε2 |
|  | 100 mg | 5,0 μg | Ε2 |
|  | 200 mg | 6,5 μg | Ε2 |
|  | 500 mg | 8,0 μg | Ε2 |
|  | 1 g | 10 μg | Ε2 |
|  | 2 g | 13 μg | Ε2 |
|  | 5 g | 16 μg | Ε2 |
|  | 10 g | 20 μg | Ε2 |
|  | 20 g | 26 μg | Ε2 |
|  | 50 g | 33 μg | Ε2 |
|  | 100 g | 53 μg | Ε2 |
|  | 200 g | 0,10 mg | Ε2 |
|  | 500 g | 0,26 mg | Ε2 |
|  | 1 kg | 0,53 mg | Ε2 |
|  | 2 kg | 1,0 mg | Ε2 |
|  | 5 kg | 2,6 mg | Ε2 |
|  | 10 kg | 10 mg | F1 |
|  | 20 kg | 41 mg | F2 |
| Non-automatic weighing instruments, Electronic / Mechanical | 1 mg | 0,6 μg | According to guide: Euramet cg18 v4.0:2015 Available standard weights OIML Class: E1 in range 1 mg ... 5 kg Total maximum available load 11,1 kg E2 in range 1 mg ... 5 kg Total maximum available load 11,7 kg F1 in range 500 g ... 10 kg Total maximum available load 30,5 kg F2 in range 1 g ... 5 kg Total maximum available load 11,1 kg M1 in range 5 kg ... 1000 kg Total maximum available load 1030 kg Calibration is performed on-site |
|  | 2 mg | 0,6 μg |  |
|  | 5 mg | 0,6 μg |  |
|  | 10 mg | 0,8 μg |  |
|  | 20 mg | 1,0 μg |  |
|  | 50 mg | 1,2 μg |  |
|  | 100 mg | 1,5 μg |  |
|  | 200 mg | 2,0 μg |  |
|  | 500 mg | 2,5 μg |  |
|  | 1 g | 3,0 μg |  |
|  | 2 g | 4,0 μg |  |
|  | 5 g | 5,0 μg |  |
|  | 10 g | 16 μg |  |
|  | 20 g | 25 μg |  |
|  | 50 g | 30 μg |  |
|  | 100 g | 50 μg |  |
|  | 200 g | 0,10 mg |  |
|  | 500 g | 0,25 mg |  |
|  | 1 kg | 0,50 mg |  |
|  | 2 kg | 1,0 mg |  |
|  | 5 kg | 2,5 mg |  |
|  | 10 kg | 5,0 mg |  |
|  | 20 kg | 21 mg |  |
|  | (20 to 1000] kg | 6 x 10-5 |  |
|  | (1000 to 10000] kg | Ι: load indication in kg d: resolution in kg n: number of substitute weights plus the reference mass mcl | Using substitute weights, each approximately equal to reference mass mcl Calibration is performed on-site. |
| Force Measurements |
| Compression testing machines | [0 to 0,5] kN | 3 mN + 0,012 % | Machines of class 0,5 |
|  | (0,5 to 1] kN | 0,24 % | ,, 1 |
|  | (1 to 10] kN | 0,08 % | ,, 0,5 |
|  | (10 to 50] kN | 0,05 % | ,, 0,5 |
|  | (50 to 120] kN | 0,13 % | ,, 0,5 |
|  | (120 to 300] kN | 0,07 % | ,, 0,5 |
|  | (300 to 700] kN | 0,05 % | ,, 0,5 |
|  | (700 to 2000] kN | 0,07 % | ,, 0,5 |
|  | (2000 to 3000] kN | 0,06 % | ,, 0,5 |
| Tensile testing machines | [0 to 0,5] kN | 3 mN + 0,012 % | ,, 0,5 |
|  | (0,5 to 5] kN | 0,07 % | ,, 0,5 |
|  | (5 to 10] kN | 0,15 % | ,, 0,5 |
|  | (10 to 15] kN | 0,11 % | ,, 0,5 |
|  | (15 to 50] kN | 0,07 % | ,, 0,5 |
|  | (50 to 120] kN | 0,17 % | ,, 0,5 |
|  | (120 to 300] kN | 0,08 % | ,, 0,5 |
|  | (300 to 700] kN | 0,06 % | ,, 0,5 |
|  |  |  | According to standards: ISO 7500-1:2018, ΕΝ 12390-4:2019 Calibration is performed on-site |
| Concrete compression testing machines: |  |  | According to standard: ΕΝ 12390-4:2019 Calibration is performed on-site. |
| Force transfer | [200 to 2000] kN | 0,01 |  |
| Alignment of upper plate symmetry axis with machine load axis | [200 to 2000] kN | 0,01 |  |
| Machine load rate | [1 to 100] kN/s | (0,012 + 0,5 x Resolution) kN/s |  |
| Machine plate hardness | [27 to 66] HRC | 1 HRC |  |
| Machine plate flatness | - | 4 μm |  |
| Machine plate parallelism | - | 6 μm |  |
| Machine plate roughness Ra | [0,2 to 5,8] μm | (0,2 + 0,05 x L) μm, L: μm |  |
| Compression testing machines fitted with ring gauges and linear displacement dial indicators | [0 to 0,5] kN | 3 mN + 0,012 % | According to standard: BS 1377-1:2016 Calibration is performed on-site. |
|  | (0,5 to 1] kN | 0,24 % |  |
|  | (1 to 10] kN | 0,08 % |  |
|  | (10 to 50] kN | 0,05 % |  |
|  | (50 to 120] kN | 0,13 % |  |
|  | (120 to 300] kN | 0,07 % |  |
|  | (300 to 700] kN | 0,05 % |  |
| Tensile testing machines fitted with ring gauges and linear displacement dial indicators | [0 to 0,5] kN | 3 mN + 0,012 % |  |
|  | (0,5 to 5] kN | 0,07 % |  |
|  | (5 to 10] kN | 0,15 % |  |
|  | (10 to 15] kN | 0,11 % |  |
|  | (15 to 50] kN | 0,07 % |  |
|  | (50 to 120] kN | 0,17 % |  |
|  | (120 to 300] kN | 0,08 % |  |
|  | (300 to 700] kN | 0,06 % |  |
| Hardness Measurements |
| Hardness testers Rockwell C | [27 to 66] HRC | 0,5 HRC | Internal method HLCP-01 using standard hardness blocks. Calibration can also be performed on-site. |
| Hardness testers Leeb D | [550 to 850] HLD | 9 HLD |  |
| Torque Measurements |
| Torque wrenches, Torque tools, Torque meters | [0,4 to 1] Nm | 1,5 % | According to standards: ISO 6789-1:2017 ISO 6789-2:2017 using torque calibrators Calibration can also be performed on-site. |
|  | (1 to 2] Nm | 0,8 % |  |
|  | (2 to 1100] Nm | 0,7 % |  |
| Impact Measurements |
| Pendulum impact testers, Impact test systems | [0,2 to 900] J Potential energy Absorbed energy indication | Force: 0,12 % Pendulum length: 0,2 mm Leveling: 0,2/1000 Angle: 0,03° Time: 0,02 s Length ~100 mm: 0,005 mm 0,12 % 0,12 % | Direct method Measurement uncertainty is calculated for: Position of centre of percussion, Potential energy, Absorbed energy indication error Calibration according to standards: Metals: ISO 148-2:2016, ASTM E23-23a, DIN 51222:2017-2, GOST 9454-78, Plastics: ISO 13802:2015, ASTM D6110-18, ISO 179-1:2023, ASTM D256-23e1, ISO 180:2023, ASTM D1822-21, ISO 8256:2023 ISO 9854-1:2023, ISO 9854-2:2023 ISO 7628:2010 |
|  | Absorbed energy [10 to 250] J | 1,5 · UCRM UCRM : reference material uncertainty | Indirect method: Calibration according to standards: Metals: ISO 148-2:2016 ASTM E23-23a |
|  |  |  | Calibration is performed on-site. |
| Temperature Measurements |
| Dry block calibrators | [-80 to 0) °C | 0,12 °C | According to guide: Euramet cg13 v4.0:2017 |
|  | [0 to 270] °C | 0,045 °C |  |
|  | (270 to 1100] °C | 1,6 °C |  |
| Self-indicating thermometers with analog or digital reading / Analog or digital temperature data loggers Self-indicating thermometers with analog or digital reading / Analog or digital temperature data loggers | [-40 to 0) °C | 0,24 °C | Internal method TLCP-03 using bath and platinum resistance thermometer. Devices under calibration must be equipped with external sensor. |
|  | [0 to 140] °C | 0,19 °C |  |
|  | [-40 to 0) °C | 0,12 °C | Internal method TLCP-03 using temperature calibrator and platinum resistance thermometer. Devices under calibration must be equipped with external sensor. Calibration can also be performed on-site. |
|  | [0 to 140] °C | 0,08 °C |  |
|  | (140 to 270] °C | 0,33 °C |  |
|  | [250 to 1100] °C | 2,0 °C | Internal method TLCP-03 using temperature calibrator and type R thermocouple. Devices under calibration must be equipped with external sensor. Calibration can also be performed on-site. |
|  | [-40 to 0) °C | 0,23 °C | Internal method TLCP-03 using temperature calibrator. Devices under calibration must be equipped with external sensor. Calibration can also be performed on-site. |
|  | [0 to 140] °C | 0,11 °C |  |
|  | [-40 to 0) °C | 0,19 °C | Internal method TLCP-04 using environmental chamber. |
|  | [0 to 75) °C | 0,12 °C |  |
|  | [75 to 100] °C | 0,14 °C |  |
| Climatic chambers, Furnaces, Ovens, Incubators, Autoclaves, Refrigerators, of any volume | [-80 to -40] °C | 0,15 °C | Using platinum resistance thermometer, positioned at one point within the chamber. |
|  | (-40 to 110] °C | 0,20 °C | Using platinum resistance thermometer, covering the volume of the chamber. |
|  | [-40 to 140] °C | 0,30 °C | Using type K thermocouples, covering the volume of the chamber. |
|  | (140 to 250] °C | 0,45 °C |  |
|  | (250 to 350] °C | 1,2 °C |  |
|  | (350 to 550] °C | 2,5 °C | Using type K thermocouples, positioned at one point within the chamber. |
|  | [-30 to 0) °C | 0,32 °C | Using temperature data loggers, covering the volume of the chamber. |
|  | [0 to 70] °C | 0,30 °C |  |
|  |  |  | Calibration according to standards: Euramet cg20 v5.0:2017 DKD-R 5-7:2018 Calibration can also be performed on-site. |
| Furnaces, Ovens | [550 to 1200] °C | 2,5 °C | Internal method TLCP-02 based on guide DKD-R 5-7:2018 using type K thermocouples. Calibration is performed at one point within the chamber. Calibration can also be performed on-site. |
| Baths | [-80 to 0) °C | 0,040 °C | Internal method TLCP-02 based on guide DKD-R 5-7:2018 using platinum resistance thermometer. Calibration can also be performed on-site. |
|  | [0 to 270] °C | 0,037 °C |  |
| Humidity Measurements |
| Environmental chambers | [20 to 50] %rh at 20, 25, 30, 40 °C | 1,0 %rh | According to guide: Euramet cg20 v.5.0:2017 Calibration is performed on-site. |
|  | (50 to 80] %rh at 20, 25, 30, 40 °C | 1,2 %rh |  |
|  | (80 to 98] %rh at 20, 25, 30, 40 °C | 1,5 %rh |  |
| Concrete specimen conditioning chambers | [20 to 98] %rh at 20 °C | 1,5 %rh | According to guide: Euramet cg20 v.5.0:2017 at one point within chamber Calibration is performed on-site. |
| Analog and digital relative humidity measuring instruments, Relative humidity recorders | [20 to 50] %rh at 20, 25, 30, 40 °C | 1,0 %rh | Internal method HRLCP-02 using environmental chamber and humidity sensor. |
|  | (50 to 80] %rh at 20, 25, 30, 40 °C | 1,6 %rh |  |
|  | (80 to 98] %rh at 20, 25, 30, 40 °C | 1,8 %rh |  |
| Volume Measurements |
| Piston-operated pipettes | [0,1 to 1] μL | 0,008 μL + 1 % | According to standard: ISO 8655-6:2022 |
|  | (1 to 2,5] μL | 0,008 μL + 0,6 % |  |
|  | (2,5 to 10] μL | 0,008 μL + 0,3 % |  |
|  | (10 to 25] μL | 0,010 μL + 0,15 % |  |
|  | (25 to 1000] μL | 0,12 % |  |
|  | (1 to 2,5] mL | 0,12 % |  |
|  | (2,5 to 5] mL | 0,12 % |  |
|  | (5 to 10] mL | 0,12 % |  |
| Piston-operated dispensers | [0,1 to 1] μL | 0,008 μL + 1 % |  |
|  | (1 to 2,5] μL | 0,008 μL + 0,6 % |  |
|  | (2,5 to 10] μL | 0,008 μL + 0,3 % |  |
|  | (10 to 25] μL | 0,010 μL + 0,15 % |  |
|  | (25 to 250] μL | 0,12 % |  |
|  | (250 to 2000] μL | 0,12 % |  |
|  | (2 to 10] mL | 0,12 % |  |
|  | (10 to 200] mL | 0,12 % |  |
| Piston-operated burettes & titrators | [0,1 to 10] mL | 0,12 % |  |
|  | (10 to 50] mL | 0,12 % |  |
|  | (50 to 100] mL | 0,12 % |  |
| Piston-operated dilutors | [0,5 to 1] μL | 0,008 μL + 1 % |  |
|  | (1 to 2,5] μL | 0,008 μL + 0,6 % |  |
|  | (2,5 to 10] μL | 0,008 μL + 0,3 % |  |
|  | (10 to 25] μL | 0,010 μL + 0,15 % |  |
|  | (25 to 250] μL | 0,12 % |  |
|  | (250 to 2000] μL | 0,12 % |  |
|  | (2 to 10] mL | 0,12 % |  |
|  | (10 to 100] mL | 0,12 % |  |
| Volumetric flasks | [1 to 5] mL | 0,50 μL + 0,04 % | According to guide: Euramet cg19 v.3.0:2018 |
|  | (5 to 50] mL | 0,04 % |  |
|  | (50 to 100] mL | 0,02 % |  |
|  | (100 to 10000] mL | 0,01 % |  |
| Graduated cylinders | [0,5 to 10] mL | 1 μL + 0,06 % |  |
|  | (10 to 50] mL | 0,04 % |  |
|  | (50 to 5000] mL | 0,03 % |  |
| Glass pipettes and burettes, syringes | [0,1 to 1] μL | 0,008 μL + 1 % | According to guide Euramet cg19 v.3.0:2018 |
|  | (1 to 2,5] μL | 0,008 μL + 0,6 % |  |
|  | (2,5 to 10] μL | 0,008 μL + 0,3 % |  |
|  | (10 to 25] μL | 0,010 μL + 0,15 % |  |
|  | (25 to 500] μL | 0,10 % |  |
|  | (0,5 to 5] mL | 0,06 % |  |
|  | (5 to 10] mL | 0,04 % |  |
|  | (10 to 100] mL | 0,02 % |  |
|  | (100 to 2000] mL | 0,01 % |  |
| Standard volume vessels | [0,1 to 1] L | 0,05 % |  |
|  | (1 to 5] L | 0,03 % |  |
|  | (5 to 20] L | 0,01 % |  |
| Pressure Measurements |
| Relative pressure gauges (gas medium) of direct reading, analog / digital | [0 kPa to 2 MPa] | (0,0003 + 0,0002 x P) MPa P: in MPa | According to guide: Euramet cg17 v.4.1:2022 Calibration can also be performed on-site. |
|  | (2 to 6] MPa | (0,015 + 0,0025 x P) MPa P: in MPa |  |
| Relative pressure gauges (liquid medium) of direct reading, analog / digital | [0 kPa to 2 MPa] | (0,0003 + 0,0002 x P) MPa P: in MPa |  |
|  | (2 to 70] MPa | (0,015 + 0,0025 x P) MPa P: in MPa |  |
| Relative pressure gauges (gas medium) of direct reading, analog / digital | [0,005 to 2] MPa abs | (0,0003 + 0,0002 x P) MPa P: in MPa |  |
|  | (2 to 6] MPa abs | (0,015 + 0,0025 x P) MPa P: in MPa |  |
| Negative relative pressure gauges (gas medium) of direct reading, analog / digital | [-95 to 0] kPa | (0,0003 + 0,0002 x P) MPa P: in MPa |  |
| Dimensional Measurements |
| Calipers, analog / digital | [0 to 300] mm |  | Internal method DLCP-01 based on guide: VDI/VDE/DGQ 2618, part 9.1:2006 |
|  | d: 0,005 mm | (3 + 0,02 x L) μm, L: mm |  |
|  | d: 0,01 mm | (7 + 0,02 x L) μm, L: mm |  |
|  | d: 0,02 mm | (10 + 0,02 x L) μm, L: mm |  |
|  | d: 0,05 mm | (24 + 0,02 x L) μm, L: mm |  |
|  | d: 0,1 mm | (48 + 0,01 x L) μm, L: mm |  |
|  | (300 to 1000] mm |  |  |
|  | d: 0,01 mm | (7 + 0,03 x L) μm, L: mm |  |
|  | d: 0,02 mm | (10 + 0,03 x L) μm, L: mm |  |
|  | d: 0,05 mm | (24 + 0,03 x L) μm, L: mm |  |
|  | d: 0,1 mm | (48 + 0,02 x L) μm, L: mm |  |
| Depth calipers, analog / digital | [0 to 300] mm |  | Internal method DLCP-01 using gauge blocks. |
|  | d: 0,005 mm | (3 + 0,02 x L) μm, L: mm |  |
|  | d: 0,01 mm | (7 + 0,02 x L) μm, L: mm |  |
|  | d: 0,02 mm | (10 + 0,02 x L) μm, L: mm |  |
|  | d: 0,05 mm | (24 + 0,02 x L) μm, L: mm |  |
|  | d: 0,1 mm | (48 + 0,01 x L) μm, L: mm |  |
| External micrometers, analog / digital | [0 to 25] mm |  | Internal method DLCP-07 based on guide: VDI/VDE/DGQ 2618, part 10.1:2001 |
|  | d: 0,0001 mm | 0,8 μm |  |
|  | [0 to 300] mm |  |  |
|  | d: 0,001 mm | (1,0 + 0,02 x L) μm, L: mm |  |
|  | d: 0,01 mm | (4,8 + 0,02 x L) μm, L: mm |  |
|  | (300 to 1000] mm |  |  |
|  | d: 0,001 mm | (1,0 + 0,02 x L) μm, L: mm |  |
|  | d: 0,01 mm | (4,8 + 0,02 x L) μm, L: mm |  |
| Micrometer heads, analog | [0 to 50] mm |  | Internal method DLCP-05 using gauge blocks & linear displacement sensor. |
|  | d: 0,001 mm | (0,45 + 0,008 x L)μm, L: mm |  |
|  | d: 0,002 mm | (0,6 + 0,006 x L)μm, L: mm |  |
|  | d: 0,005 mm | (0,92 + 0,007 x L)μm, L: mm |  |
|  | d: 0,01 mm | (1,6 + 0,004 x L)μm, L: mm |  |
| Micrometer heads, digital | [0 to 50] mm |  |  |
|  | d: 0,001 mm | (0,92 + 0,007 x L)μm, L: mm |  |
|  | d: 0,002 mm | (1,6 + 0,004 x L)μm, L: mm |  |
|  | d: 0,005 mm | (3,9 + 0,002 x L)μm, L: mm |  |
|  | d: 0,01 mm | (7,8 + 0,001 x L)μm, L: mm |  |
| Digital indicators, digital Extensometers, LVDT, of lever / plung type | [0 to 100] mm |  | Internal method DLCP-02 using micrometer head. |
|  | d: 0,001 mm | (1,2 + 0,05 x L) μm, L: mm |  |
|  | d: 0,002 mm | (1,9 + 0,04 x L) μm, L: mm |  |
|  | d: 0,005 mm | (4,0 + 0,04 x L) μm, L: mm |  |
|  | d: 0,01 mm | (7,8 + 0,03 x L) μm, L: mm |  |
|  | d: 0,02 mm | (16 + 0,02 x L) μm, L: mm |  |
|  | d: 0,05 mm | (39 + 0,01 x L) μm, L: mm |  |
|  | d: 0,1 mm | (78 + 0,01 x L) μm, L: mm |  |
|  | [0 to 100] mm |  | Internal method DLCP-02 using micrometer head. Calibration is performed on-site. |
|  | d: 0,001 mm | (1,2 + 0,06 x L) μm, L: mm |  |
|  | d: 0,002 mm | (1,9 + 0,06 x L) μm, L: mm |  |
|  | d: 0,005 mm | (4,0 + 0,04 x L) μm, L: mm |  |
|  | d: 0,01 mm | (7,8 + 0,03 x L) μm, L: mm |  |
|  | d: 0,02 mm | (16 + 0,02 x L) μm, L: mm |  |
|  | d: 0,05 mm | (39 + 0,01 x L) μm, L: mm |  |
|  | d: 0,1 mm | (78 + 0,01 x L) μm, L: mm |  |
|  | [0 to 10] mm |  | Internal method DLCP-02 using gauge blocks. |
|  | d: 0,0001 mm | (0,22 + 0,013 x L)μm, L: mm |  |
|  | d: 0,0002 mm | (0,26 + 0,011 x L)μm, L: mm |  |
|  | d: 0,0005 mm | (0,46 + 0,007 x L)μm, L: mm |  |
| Digital indicators, digital Extensometers, LVDT, of lever / plung type | [0 to 100] mm |  | Internal method DLCP-02 using gauge blocks. |
|  | d: 0,001 mm | (0,8 + 0,02 x L) μm, L: mm |  |
|  | d: 0,002 mm | (1,6 + 0,02 x L) μm, L: mm |  |
|  | d: 0,005 mm | (3,9 + 0,01 x L) μm, L: mm |  |
|  | d: 0,01 mm | (7,8 + 0,01 x L) μm, L: mm |  |
|  | d: 0,02 mm | (16 + 0,01 x L) μm, L: mm |  |
|  | d: 0,05 mm | (39 + 0,01 x L) μm, L: mm |  |
|  | d: 0,1 mm | (78 + 0,01 x L) μm, L: mm |  |
|  | [0 to 100] mm |  | Internal method DLCP-02 using gauge blocks. Calibration is performed on-site. |
|  | d: 0,001 mm | (0,8 + 0,04 x L) μm, L: mm |  |
|  | d: 0,002 mm | (1,6 + 0,04 x L) μm, L: mm |  |
|  | d: 0,005 mm | (3,9 + 0,03 x L) μm, L: mm |  |
|  | d: 0,01 mm | (7,8 + 0,02 x L) μm, L: mm |  |
|  | d: 0,02 mm | (16 + 0,01 x L) μm, L: mm |  |
|  | d: 0,05 mm | (39 + 0,01 x L) μm, L: mm |  |
|  | d: 0,1 mm | (78 + 0,01 x L) μm, L: mm |  |
| Dial indicators, dial Extensometers, LVDT, of lever / plug type | [0 to 100] mm |  | Internal method DLCP-02 using micrometer head. |
|  | d: 0,001 mm | (0,9 + 0,04 x L) μm, L: mm |  |
|  | d: 0,002 mm | (1,1 + 0,05 x L) μm, L: mm |  |
|  | d: 0,005 mm | (1,8 + 0,04 x L) μm, L: mm |  |
|  | d: 0,01 mm | (3,3 + 0,03 x L) μm, L: mm |  |
|  | d: 0,02 mm | (6,0 + 0,03 x L) μm, L: mm |  |
|  | d: 0,05 mm | (14 + 0,03 x L) μm, L: mm |  |
|  | d: 0,1 mm | (28 + 0,01 x L) μm, L: mm |  |
|  | [0 to 100] mm |  | Internal method DLCP-02 using micrometer head. Calibration is performed on-site. |
|  | d: 0,001 mm | (0,9 + 0,11 x L) μm, L: mm |  |
|  | d: 0,002 mm | (1,1 + 0,06 x L) μm, L: mm |  |
|  | d: 0,005 mm | (1,8 + 0,06 x L) μm, L: mm |  |
|  | d: 0,01 mm | (3,3 + 0,05 x L) μm, L: mm |  |
|  | d: 0,02 mm | (6,0 + 0,04 x L) μm, L: mm |  |
|  | d: 0,05 mm | (14 + 0,04 x L) μm, L: mm |  |
|  | d: 0,1 mm | (28 + 0,01 x L) μm, L: mm |  |
|  | [0 to 100] mm |  | Internal method DLCP-02 using gauge blocks. |
|  | d: 0,001 mm | (0,33 + 0,02 x L) μm, L: mm |  |
|  | d: 0,002 mm | (0,6 + 0,02 x L) μm, L: mm |  |
|  | d: 0,005 mm | (1,4 + 0,02 x L) μm, L: mm |  |
|  | d: 0,01 mm | (2,8 + 0,02 x L) μm, L: mm |  |
|  | d: 0,02 mm | (5,6 + 0,02 x L) μm, L: mm |  |
|  | d: 0,05 mm | (12 + 0,01 x L) μm, L: mm |  |
|  | d: 0,1 mm | (28 + 0,01 x L) μm, L: mm |  |
| Dial indicators, dial extensiometers, LVDT, of lever / plug type | [0 to 100] mm |  | Internal method DLCP-02 using gauge blocks. Calibration is performed on-site. |
|  | d: 0,001 mm | (0,33 + 0,04 x L) μm, L: mm |  |
|  | d: 0,002 mm | (0,6 + 0,04 x L) μm, L: mm |  |
|  | d: 0,005 mm | (1,4 + 0,04 x L) μm, L: mm |  |
|  | d: 0,01 mm | (2,8 + 0,03 x L) μm, L: mm |  |
|  | d: 0,02 mm | (5,6 + 0,03 x L) μm, L: mm |  |
|  | d: 0,05 mm | (12 + 0,02 x L) μm, L: mm |  |
|  | d: 0,1 mm | (28 + 0,01 x L) μm, L: mm |  |
| Sieves: wire cloth, perforated plate with rectangular / circular holes, bar sieves | Nominal opening |  | According to standards: ISO 3310-1:2016 ISO 3310-2:2013 ASTM E11:2022 ISO 5223:1995 EN 933-3:2012 |
|  | [0,02 to 0,5] mm | 1,4 μm |  |
|  | (0,5 to 1,0] mm | 1,6 μm |  |
|  | (1 to 30] mm | 2,5 μm |  |
|  | (30 to 50] mm | 4,1 μm |  |
|  | (50 to 125] mm | 12 μm |  |
| Profile projectors, micrometers with scales, optical measuring instruments with scales, analog / digital (axis and angle measurements) | [0 to 5] mm magnification ≥ x30 |  | Internal method DLCP-03 using glass microscales. Calibration can also be performed on-site. |
|  | d: 0,0005 mm | (0,69 + 0,004 x L)μm, L: mm |  |
|  | d: 0,001 mm | (1,3 + 0,003 x L)μm, L: mm |  |
|  | d: 0,002 mm | (2,5 + 0,001 x L)μm, L: mm |  |
|  | d: 0,005 mm | (3,9 + 0,001 x L)μm, L: mm |  |
|  | d: 0,01 mm | (7,8 + 0,001 x L)μm, L: mm |  |
|  | (5 to 50] mm magnification ≥ x10 |  |  |
|  | d: 0,0005 mm | (0,89 + 0,019 x L)μm, L: mm |  |
|  | d: 0,001 mm | (1,3 + 0,016 x L)μm, L: mm |  |
|  | d: 0,002 mm | (2,6 + 0,010 x L)μm, L: mm |  |
|  | d: 0,005 mm | (4,0 + 0,011 x L)μm, L: mm |  |
|  | d: 0,01 mm | (7,8 + 0,003 x L)μm, L: mm |  |
|  | (50 to 300] mm magnification ≥ x3 |  |  |
|  | d: 0,0005 mm | (2,2 + 0,030 x L) μm, L: mm |  |
|  | d: 0,001 mm | (2,1 + 0,031 x L) μm, L: mm |  |
|  | d: 0,002 mm | (3,1 + 0,029 x L) μm, L: mm |  |
|  | d: 0,005 mm | (4,6 + 0,025 x L) μm, L: mm |  |
|  | d: 0,01 mm | (8,0 + 0,020 x L) μm, L: mm |  |
|  | (300 to 1000] mm magnification ≥ 30 |  |  |
|  | d: 0,001 mm | (16 + 0,014 x L) μm, L: mm |  |
|  | [0° to 360°] | 1' |  |
| Orthogonal moulds |  |  | According to standard: ΕΝ 12390-1: 2021 |
| - External dimensions | (0 to 300] mm | 0,02 mm |  |
| - Internal dimensions | (0 to 300] mm | 0,02 mm |  |
| - Depth dimensions | (0 to 300] mm | 0,02 mm |  |
| - Perpendicularity | - | 5 μm |  |
| - Flatness | - | 4 μm |  |
| Cylindrical moulds |  |  | According to standards: ΕΝ 12390-1:2021 EN 13286-2:2010 ASTM D698-12, ASTM D1557-12 ΕΝ 12697-30:2018 ASTM D6926-20, ASTM D5581-07a ΕΝ 13286-4:2021, ASTM D1883-21 |
| - External dimensions | (0 to 300] mm | 0,02 mm |  |
| - Internal dimensions | (0 to 300] mm | 0,02 mm |  |
| - Depth dimensions | (0 to 300] mm | 0,02 mm |  |
| - Perpendicularity | - | 5 μm |  |
| - Flatness | - | 4 μm |  |
| - Spacer - dimensions | (0 to 300] mm | 0,02 mm |  |
| Graduated rules, tank level bars | [0 to 3] m | (N x 0,07) mm N: length in integral m | Internal method DLCP-10 based on guide: OIML R35-1:2007, Amendment 2014 Calibration in 1 m steps. |
| Measuring tapes, circometers | [0 to 200] m |  |  |
| Extensometers | [0 to 50] mm | (0,3 + 0,8 x L) μm L: mm | According to standards: ISO 9513:2012 ASTM E83-23 Calibration can also be performed on-site. |
| Electrical measurements |
| Voltage DC (Generating) / Voltmeters DC | [0,1 to 20] mV | 0,004 mV + 100 x 10-6 V | According to guide: Euramet cg-15 v.3.0:2015 using standard multifunction calibrator. |
|  | (20 to 200] mV | 0,006 mV + 30 x 10-6 V |  |
|  | (0,2 to2] V | 0,020 mV + 15 x 10-6 V |  |
|  | (2 to 20] V | 0,15 mV + 15 x 10-6 V |  |
|  | (20 to 200] V | 6,0 mV + 30 x 10-6 V |  |
|  | (200 to 1000] V | 30 mV + 50 x 10-6 V |  |
| Voltage DC (Measuring) / Voltage sources DC | [0,001 to 100] mV | 0,94 μV + 6 x 10-6 V | Internal method ELCP-04 using standard multimeter. |
|  | (0,1 to 1] V | 3,1 μV + 6,1 x 10-6 V |  |
|  | (1 to 10] V | 16 μV + 7 x 10-6 V |  |
|  | (10 to 100] V | 0,25 mV + 8,3 x 10-6 V |  |
|  | (100 to 1000] V | 3,5 mV + 7,9 x 10-6 V |  |
|  | [0,001 to 100] mV | 13 μV | Internal method ELCP-04 using standard multimeter. Calibration can also be performed on-site. |
|  | (0,1 to 1] V | 53 μV |  |
|  | (1 to 10] V | 0,45 mV |  |
|  | (10 to 100] V | 13 mV |  |
|  | (100 to 1000] V | 0,12 V |  |
| Voltage AC (Generating) / Voltmeters AC | [0,1 to 20] mV |  | According to guide: Euramet cg-15 v.3.0:2015 using standard multifunction calibrator. |
|  | [10 to 45] Hz | 0,25 mV + 500 x 10-6 V |  |
|  | (45 to 1000] Hz | 0,10 mV + 500 x 10-6 V |  |
|  | (1 to 10] kHz | 0,15 mV + 500 x 10-6 V |  |
|  | (20 to 200] mV |  |  |
|  | [10 to 45] Hz | 0,25 mV + 500 x 10-6 V |  |
|  | (45 to 1000] Hz | 0,10 mV + 400 x 10-6 V |  |
|  | (1 to 10] kHz | 0,15 mV + 400 x 10-6 V |  |
|  | (0,2 to 2] V |  |  |
|  | [10 to 45] Hz | 0,50 mV + 800 x 10-6 V |  |
|  | (45 to 1000] Hz | 0,17 mV + 300 x 10-6 V |  |
|  | (1 to 10] kHz | 0,25 mV + 300 x 10-6 V |  |
|  | (2 to 20] V |  |  |
|  | [10 to 45] Hz | 4,0 mV + 800 x 10-6 V |  |
|  | (45 to 1000] Hz | 2,0 mV + 300 x 10-6 V |  |
|  | (1 to 10] kHz | 3,0 mV + 300 x 10-6 V |  |
|  | (20 to 200] V | 20 mV + 600 x 10-6 V |  |
|  | [40 to 1000] Hz |  |  |
|  | (200 to 1000] V | 91 mV + 800 x 10-6 V |  |
|  | [40 to 1000] Hz |  |  |
| Voltage AC (Measuring) / Voltage sources AC | [0,001 to 10] mV | 5,1 μV + 83 x 10-6 V | Internal method ELCP-04 using standard multimeter. |
|  | [45 to 1000] Hz |  |  |
|  | (10 to 100] mV | 21 μV + 20 x 10-6 V |  |
|  | [45 to 1000] Hz |  |  |
|  | (0,1 to 1] V | 0,045 mV + 56 x 10-6 V |  |
|  | [45 to 1000] Hz |  |  |
|  | (1 to 10] V | 0,22 mV + 69 x 10-6 V |  |
|  | [45 to 1000] Hz |  |  |
|  | (10 to 100] V | 4 mV + 190 x 10-6 V |  |
|  | [45 to 1000] Hz |  |  |
|  | (100 to 1000] V | 32 mV + 390 x 10-6 V |  |
|  | [45 to 1000] Hz |  |  |
|  | [10 to 100] mV | 0,21 mV | Internal method ELCP-04 using standard multimeter. Calibration can also be performed on-site. |
|  | [45 to 1000] Hz |  |  |
|  | (0,1 to 1] V | 0,67 mV |  |
|  | [45 to 1000] Hz |  |  |
|  | (1 to 10] V | 7,2 mV |  |
|  | [45 to 1000] Hz |  |  |
|  | (10 to 100] V | 0,11 V |  |
|  | [45 to 1000] Hz |  |  |
|  | (100 to 1000] V | 0,96 V |  |
|  | [45 to 1000] Hz |  |  |
| Current DC (Generating) / Ammeters DC | [1 to 200] μA | 15 nA + 150 x 10-6 I | According to guide: Euramet cg-15 v.3.0:2015 using standard multifunction calibrator. |
|  | (0,2 to 2] mA | 42 nA + 100 x 10-6 I |  |
|  | (2 to 20] mA | 0,22 μA + 80 x 10-6 Ι |  |
|  | (20 to 200] mA | 3,4 μA + 80 x 10-6 I |  |
|  | (0,2 to 2] A | 43 μA + 250 x 10-6 I |  |
|  | (2 to 20] A | 2,1 mA + 600 x 10-6 I |  |
| Current DC (Generating) / Clamp meters DC | [1 to 100] A | 0,05 A + 5000 x 10-6 I | Internal method ELCP-12 using standard current source. |
|  | (100 to 1000] A | 0,36 A + 4800 x 10-6 I |  |
| Current DC (Measuring) / Current sources DC | [1 to 100] nA | 6 nA + 60 x 10-6 I | Internal method ELCP-05 using standard multimeter. |
|  | (0,1 to 1] μA | 6 nA + 20 x 10-6 I |  |
|  | (1 to 10] μA | 6,3 nA + 20 x 10-6 I |  |
|  | (10 to 100] μA | 9,2 nA + 4,5 x 10-6 I |  |
|  | (0,1 to 1] mA | 33 nA + 9 x 10-6 I |  |
|  | (1 to 10] mA | 0,3 μA + 9,4 x 10-6 I |  |
|  | (10 to 100] mA | 4,1 μA + 18 x 10-6 I |  |
|  | (0,1 to 1] A | 0,055 mA + 85 x 10-6 I |  |
|  | [0,1 to 100] μA | 43 nA + 340 x 10-6 I | Internal method ELCP-05 using standard multimeter. Calibration can also be performed on-site. |
|  | (0,1 to 1] mA | 0,18 μA + 420 x 10-6 I |  |
|  | (1 to 10] mA | 1,9 μA + 470 x 10-6 I |  |
|  | (10 to 100] mA | 14 μA + 440 x 10-6 I |  |
|  | (0,1 to 1] A | 0,36 mA + 840 x 10-6 I |  |
|  | (1 to 10] A | 8,4 mA + 1900 x 10-6 I |  |
|  | [1 to 10] A | 32 μA + 99 x 10-6 I | Internal method ELCP-23 using standard ammeter. |
|  | (10 to 20] A | 100 μA + 99 x 10-6 I |  |
|  | (20 to 100] A | 180 μA + 200 x 10-6 I |  |
|  | (100 to 250] A | 490 μA + 200 x 10-6 I |  |
|  | [1 to 10] A | 1,4 mA + 160 x 10-6 I | Internal method ELCP-23 using standard ammeter. Calibration can also be performed on-site. |
|  | (10 to 20] A | 6,1 mA + 150 x 10-6 I |  |
|  | (20 to 100] A | 13 mA + 220 x 10-6 I |  |
|  | (100 to 250] A | 58 mA + 200 x 10-6 I |  |
|  | (250 to 1000] A | 81 mA + 2500 x 10-6 I |  |
| Current AC (Generating) / Ammeters AC | [1 to 200] μA | 300 nA + 700 x 10-6 I | According to guide: Euramet cg-15 v.3.0:2015 using standard multifunction calibrator. |
|  | [20 to 1000] Hz |  |  |
|  | (0,2 to 2] mA | 320 nA + 500 x 10-6 I |  |
|  | [20 to 1000] Hz |  |  |
|  | (2 to 20] mA | 3,2 μA + 500 x 10-6 I |  |
|  | [20 to 1000] Hz |  |  |
|  | (20 to 200] mA | 32 μA + 500 x 10-6 I |  |
|  | [20 to 1000] Hz |  |  |
|  | (0,2 to 2] A | 0,53 mA + 1000 x 10-6 I |  |
|  | [20 to 1000] Hz |  |  |
|  | (2 to 20] A | 5,4 mA + 2000 x 10-6 I |  |
|  | [20 to 1000] Hz |  |  |
| Current AC (Generating) / Clamp meters AC | [1 to 100] A | 0,32 A + 4800 x 10-6 I | Internal method ELCP-12 using standard current source. |
|  | [45 to 65] Hz |  |  |
|  | (100 to 1000] A | 2,0 A + 4200 x 10-6 I |  |
|  | [45 to 65] Hz |  |  |
| Current AC (Measuring) / Current sources AC | [10 to 100] μA | 34 nA + 580 x 10-6 I | Internal method ELCP-05 using standard multimeter. |
|  | [45 to 100] Hz |  |  |
|  | (0,1 to 1] mA | 0,23 μA + 580 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | (1 to 10] mA | 2,3 μA + 580 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | (10 to 100] mA | 23 μA + 580 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | (0,1 to 1] A | 0,33 mA + 700 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | [0,1 to 10] mA | 14 μA + 4700 x 10-6 I | Internal method ELCP-05 using standard multimeter. Calibration can also be performed on-site. |
|  | [45 to 100] Hz |  |  |
|  | (10 to 100] mA | 140 μA + 4700 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | (0,1 to 1] A | 2,1 mA + 4200 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | (1 to 10] A | 33 mA + 3500 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | [1 to 20] A | 520 μA + 970 x 10-6 I | Internal method ELCP-23 using standard ammeter. |
|  | [45 to 100] Hz |  |  |
|  | (20 to 250] A | 36 mA + 900 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | [1 to 10] A | 22 mA + 1500 x 10-6 I | Internal method ELCP-23 using standard ammeter. Calibration can also be performed on-site. |
|  | [45 to 100] Hz |  |  |
|  | (10 to 20] A | 110 mA + 2000 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | (20 to 100] A | 220 mA + 1600 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | (100 to 250] A | 1,1 A + 2000 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
|  | (250 to 1000] A | 4,4 A + 1900 x 10-6 I |  |
|  | [45 to 100] Hz |  |  |
| Resistance DC (Generating) / Resistance meters Resistance DC (Generating) / Resistance meters | 1 Ω | 1,8 mΩ | According to guide: Euramet cg-15 v.3.0:2015 using standard multifunction calibrator 2 or 4 wires. |
|  | 10 Ω | 1,7 mΩ |  |
|  | 100 Ω | 4,1 mΩ |  |
|  | 1 kΩ | 22 mΩ |  |
|  | 10 kΩ | 0,20 Ω |  |
|  | 100 kΩ | 3,1 Ω |  |
|  | 1 MΩ | 0,15 kΩ |  |
|  | 10 MΩ | 10 kΩ |  |
|  | 100 MΩ | 1,0 MΩ |  |
|  | 1 GΩ | 0,10 GΩ |  |
|  | [0,01 to 0,1] Ω | 0,50 mΩ | According to guide: Euramet cg-15 v.3.0:2015 using standard decade box. |
|  | [0,1 to 1] Ω | 0,53 mΩ |  |
|  | (1 to 10] Ω | 0,50 mΩ + 23 x 10-6 R |  |
|  | (10 to 100] Ω | 0,62 mΩ + 23 x 10-6 R |  |
|  | (0,1 to 1] kΩ | 0,80 mΩ + 22 x 10-6 R |  |
|  | (1 to 10] kΩ | 2,1 mΩ + 13 x 10-6 R |  |
|  | (10 to 100] kΩ | 23 mΩ + 13 x 10-6 R |  |
|  | (0,1 to 1] MΩ | 0,94 Ω + 17 x 10-6 R |  |
| Resistance / Insulation resistance meters | [1 to 10] MΩ | 1,5 x 10-3 R | Internal method ELCP-17 using standard decade box. |
|  | (10 to 100] MΩ | 1,6 x 10-3 R |  |
|  | (0,1 to 1] GΩ | -150 kΩ + 6,4 x 10-3 R |  |
|  | (1 to 10] GΩ | -7,5 ΜΩ + 13 x 10-3 R |  |
|  | (10 to 50] GΩ | 2,1 MΩ+ 14 x 10-3 R |  |
|  | (50 to 100] GΩ | 95 MΩ + 31 x 10-3 R |  |
| Resistance (4 wire Measuring) / Instruments of nominal resistance | [0,001 to 10] Ω | 0,21 mΩ + 8 x 10-6 R | Internal method ELCP-06 using standard multimeter. |
|  | (10 to 100] Ω | 0,92 mΩ + 10 x 10-6 R |  |
|  | (0,1 to 1] kΩ | 6,6 mΩ + 6,3 x 10-6 R |  |
|  | (1 to 10] kΩ | 65 mΩ + 6,4 x 10-6 R |  |
|  | (10 to 100] kΩ | 0,81 Ω + 6 x 10-6 R |  |
|  | (0,1 to 1] MΩ | 14 Ω + 9 x 10-6 R |  |
|  | (1 to 10] MΩ | 290 Ω + 39 x 10-6 R |  |
|  | (10 to 100] MΩ | 9 kΩ + 440 x 10-6 R |  |
|  | (0,1 to 1] GΩ | 0,015 MΩ + 5000 x 10-6 R |  |
| Time / Frequency measurements |
| Timers | (0 s to 20 h] | 0,07 s | Internal method ELCP-07 using standard timer. Calibration can also be performed on-site. |
| Optical tachometers | [10 to 120 000] rpm | 0,13 rpm | Internal method ELCP-09 using standard function generator and standard frequency meter. |
| Rotation frequency / Reciprocating movement frequency Centrifuges, Mixers, Shakers, Vibrating tables | (0 to 1000] rpm | 0,01 rpm + 0,1 % x f | Internal method ELCP-14 using standard optical tachometer. Calibration can also be performed on-site. |
|  | (1000 to 3000] rpm | 0,1 rpm + 0,1 % x f |  |
|  | (3000 to 60000] rpm | 1 rpm + 0,1 % x f |  |

* Where expanded uncertainty (with a probability of 95% coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration & Measurement Capability (CMC), includes the measured quantity, measurement range and measurement uncertainty and expresses the smallest measurement uncertainty that can be achieved during a calibration.

Permanent laboratory premises address: 1 Lesvou str., 171 23, N. Smirni, Athens, Greece

Approved Signatories: K. Konstantinou, D. Grigoriou, P. Damigos, A Skandalis, E. Cherouvim

This scope of Accreditation replaces the previous one dated 05.05.2022.

The Accreditation Certificate No. 900-5, to ELOT EN ISO/IEC 17025:2017, is valid until 27.01.2026.

Athens, 26.03.2024

Christos Nestoras         			 CEO of ESYD