Hellenic Accreditation System

| Annex F2/5 to the Certificate No. 1116-3 SCOPE of ACCREDITATION of the Calibration Laboratory of Hellenic Aerospace Industry (H.A.I.) S.A. |
| --- |
| Measured Quantity Instrument or Gauge | Range | Expanded Measurement Uncertainty (k=2)* | Remarks |
| Dimensional Measurements |
| Length/ Gauge blocks of orthogonal cross-section made of Chromium carbide (Croblox) | [0.5 to 25] mm | 0.10 μm | Calibration using Master Gauge Blocks & Mechanical Comparator |
|  | (25 to 50] mm | 0.14 μm |  |
|  | (50 to 75] mm | 0.19 μm |  |
|  | (75 to 100] mm | 0.23 μm |  |
| Length/ Gauge blocks of orthogonal cross-section made of Steel | [0.5 to 25] mm | 0.12 μm |  |
|  | (25 to 50] mm | 0.17 μm |  |
|  | (50 to 75] mm | 0.23 μm |  |
|  | (75 to 100] mm | 0.28 μm |  |
| Length/ Vernier Micrometers division 0.001 in or 0.0001 in | [0 to 5] in | 660 μin (division 0.001 in) 150 μin (division 0.0001 in) | Calibration using Master Gauge Blocks & Optical Flats |
| Length/ Digital Micrometers division 0.00005 in | [0 to 1] in | 40 μin |  |
|  | (1 to 4] in | 111 μin |  |
|  | (4 to 5] in | 140 μin |  |
| Length/ Digital calipers division 0.0005 in or 0.010 mm | [0 to 6] in | 0.00033 in or 0.0060 mm | Calibration using Master Gauge Blocks |
| Length/ Vernier calipers division 0.001 in or 0.025 mm |  | 0.00067 in or 0.011 mm |  |
| Length/ Dial calipers division 0.001 in |  | 0.00042 in |  |

| Measured Quantity Instrument or Gauge | Range | Expanded Measurement Uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Electrical Measurements |
| DC Voltage / DC voltage transfer standards | 1 V | 2.0 ∙ 10-6 V | Comparative calibration method using standard DC voltage source and null detector. |
|  | 1.018 V | 2.0 ∙ 10-6 V |  |
|  | 1.019 V | 2.0 ∙ 10-6 V |  |
|  | 10 V | 18 · 10-6 V |  |
| DC Voltage / DC voltage measuring instruments | [0 to 100] mV | 6.0 · 10-7 V | According to Euramet Guideline cg-15 v.3.0: 2015, using a standard multifunction calibrator. |
|  | (0.1 to 1] V | 3.5 · 10-6 V |  |
|  | (1 to 10] V | 24 · 10-6 V |  |
|  | (10 to 100] V | 500 · 10-6 V |  |
|  | (100 to 1000] V | 3.0 · 10-3 V |  |
| AC Voltage / AC voltage measuring instruments | [0 to 100] mV | / |  | According to Euramet Guideline cg-15 v.3.0: 2015, using a standard multifunction calibrator. |
|  |  | [40 to 1000] Hz | 17 · 10-6 V |  |
|  |  | (1 to 20] KHz | 16 · 10-6 V |  |
|  |  | (20 to 100] KHz | 70 · 10-6 V |  |
|  | (0.1 to 1] V | / |  |  |
|  |  | [40 to 1000] Hz | 55 · 10-6 V |  |
|  |  | (1 to 20] KHz | 54 · 10-6 V |  |
|  |  | (20 to 100] KHz | 57· 10-6 V |  |
|  | (1 to 10] V | / |  |  |
|  |  | [40 to 1000] Hz | 595 · 10-6 V |  |
|  |  | (1 to 20] KHz | 600 · 10-6 V |  |
|  |  | (20 to 100] KHz | 700 · 10-6 V |  |
|  | (10 to 100] V | / |  |  |
|  |  | [40 to 1000] Hz | 4.5 · 10-3 V |  |
|  |  | (1 to 20] KHz | 4.0 · 10-3 V |  |
|  |  | (20 to 100] KHz | 8.0 · 10-3 V |  |
|  | (100 to 600] V | / |  |  |
|  |  | [40 to 1000] Hz | 55 · 10-3 V |  |
| DC Current / DC current measuring instruments | [0.1 to 10] mA | 0.2 · 10-6 A | According to Euramet Guideline cg-15 v.3.0: 2015, using a standard multifunction calibrator. |
|  | (10 to 100] mA | 3.0 · 10-6 A |  |
|  | (0.1 to 1] A | 38 · 10-6 A |  |
|  | (1 to 2] A | 93 · 10-6 A |  |
| AC Current / AC current measuring instruments | 0 mA to 1 A | / |  | According to Euramet Guideline cg-15 v.3.0: 2015, using a standard multifunction calibrator. |
|  |  | [40 to 1000] Hz | 320 · 10-6 A |  |
| Current shunt | 0.001 Ω | 0.2 · 10-6 Ω | Comparative calibration method using standard current shunt and digital multimeter. |
|  | 0.01 Ω | 0.50 ∙ 10-6 Ω |  |
|  | 0.1 Ω | 2.8 · 10-6 Ω |  |
| Resistance/ Calibration of Resistors | 1 Ω | 15 ∙ 10-6 Ω | Comparative calibration method using standard resistance and digital multimeter. |
|  | 10 Ω | 3.0 · 10-2 Ω | Comparative calibration method using standard resistors and Ratio bridge. |
|  | 100 Ω | 3.0 · 10-2 Ω |  |
|  | 1 kΩ | 3.0 · 10-2 Ω |  |
|  | 10 kΩ | 3.0 · 10-2 Ω |  |
|  | 100 kΩ | 24 · 10-2 Ω |  |
|  | 1 MΩ | 2.6 Ω |  |
|  | 10 MΩ | 30 Ω |  |
| Measured Quantity Instrument or Gauge | Range | Expanded Measurement Uncertainty (k=2)* | Remarks |
| Resistance / Resistance Measuring instruments | [10 to 100] Ω | 1.0 · 10-3 Ω | According to Euramet Guideline cg-15 v.3.0: 2015, using a standard multifunction calibrator. |
|  | (0,1 to 1] kΩ | 7.5 · 10-3 Ω |  |
|  | (1 to 10] kΩ | 170 · 10-3 Ω |  |
|  | (10 to 100] kΩ | 2.5 Ω |  |
|  | (0,1 to 1] MΩ | 27 Ω |  |
|  | (1 to 10] MΩ | 620 Ω |  |

*	Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

Site of assessment:  Permanent laboratory premises, Schimatari, Viotia.

Approved Signatories: Κ. Vechlivanos, M. Karkaletsi.

This Scope of Accreditation replaces the previous one dated 23.05.2022.

The Accreditation Certificate No. 1116-3, to ELOT EN ISO/IEC 17025:2017, is valid until 26.02.2026.

Athens, 03.07.2023

Christos Nestoras

CEO of ESYD