Hellenic Accreditation System

Annex F2/5 to the Certificate No. 1192-2

SCOPE of ACCREDITATION

of the

Calibration Laboratory

of

TSAOUSIDIS ELEFTHERIOS

(ΖΥΓΙSIS)

| Calibration item / Parameters | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass Measurements |
|  |  |  | Calibration capability of mass standards of class OIML-R-111-1:2004 |
| Mass Standards | 100g | 15 mg | Up to (including) M3 |
|  | 200g | 30 mg | >> Μ3 |
|  | 500g | 25 mg | >> Μ2 |
|  | 1 kg | 16 mg | >> Μ1 |
|  | 2 kg | 30 mg | >> Μ1 |
|  | 5kg | 25 mg | >> F2 |
|  | 10kg | 50 mg | >> F2 |
|  | 20kg | 100 mg | >> F2 |
|  | 50kg | 250 mg | >> Μ1 |
|  | 100kg | 1,6 g | >> Μ1 |
|  | 200kg | 3,0 g | >> Μ1 |
|  | 500 kg | 25 g | >> Μ2 |
|  | 1000 kg | 50 g | >> Μ2 |
| Non-automatic weighing instruments | 100 g | 0,17 mg | According to the guide EURAMET cg-18 v4.0 (11/2015). Calibration is carried out on-site. Use of mass standards of OIML class: F1 100g to 5kg, max grouped load 10 kg F1: 10 kg F1: 20 kg F2: 20 kg (10 pieces), max grouped load: 200 kg M1: 50 kg (2 pieces), max grouped load: 100 kg M1: 100kg (2 pieces), max grouped load: 200 kg M1: 200 kg (5 pieces), max grouped load: 1000 kg M2: 500 kg (2 pieces), max grouped load: 1000 kg M2: 1000kg (10 pieces), max grouped load: 10000 kg |
|  | 200 g | 0,30 mg |  |
|  | 500 g | 0,80 mg |  |
|  | 1 kg | 1,6 mg |  |
|  | 2 kg | 3,1 mg |  |
|  | 5 kg | 8,3 mg |  |
|  | 10 kg | 16 mg |  |
|  | 20 kg | 30 mg |  |
|  | 50 kg | 0,25 g |  |
|  | 100 kg | 0,50 g |  |
|  | 200 kg | 1,0 g |  |
|  | 500 kg | 8,0 g |  |
|  | 1000 kg | 50 g |  |
|  | [1000 kg to 12000 kg] | 50 g to 600 g |  |
|  | [12000 kg to 60000 kg] | (kg) Ι: Indication of load in kg d: Resolution in kg n: Number of substitutes weights and reference mass mcl. | According to EURAMET cg-18 v4.0 (11/2015) by use of substitution loads. Calibration is carried out on-site. |

* Where uncertainty (with a probability of 95% coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

Calibration & Measurement Capability (CMC), includes the measured quantity, measurement range and measurement uncertainty and expresses the smallest measurement uncertainty that can be achieved during a calibration.

Permanent laboratory premises address: Ampelotopoi No. 715- Lakkoma, 63080, Chalkidiki.

Approved Signatories: Athanasios Kanoulas, Eleftherios Tsaousidis.

This scope of Accreditation replaces the previous one dated 18.01.2024.

The Accreditation Certificate No. 1192-2, to ELOT EN ISO/IEC 17025:2017, is valid until 20.01.2029.

Athens, 12.02.2024

Christos Nestoras

CEO of ESYD