Hellenic Accreditation System

Annex F2/8 to the Certificate No. 364-7 SCOPE of ACCREDITATION of the Calibration Laboratory BUREAU VERITAS HELLAS S.M.S.A.

| Calibration item / Parameters | Range of measurement | Expanded measurement uncertainty (k=2) * | Remarks |
| --- | --- | --- | --- |
| Volume Measurements |
| Volume / Vertical cylindrical tanks | Up to 2500 m3 Maximum tank diameter 14 m | Height Diameter 1,2m 2,5 m 5 m >7,5 m 1 m 0,43% 0,23% 0,16% 0,14% > 3 m - 0,22% 0,14% 0,12% on nominal tank volume | Calibration according to: ISO 7507-1:2003 Bottom volume measurement: ISO 4269:2001 or ISO 7507-1:2003 |
|  | Up to 140000 m3 Minimum tank diameter 8 m | Height Diameter m 8 m > 12 m 8 m 0,17 % 0,16 % 0,15 % 20 m 0,11 % 0,09 % 0,08 % > 50 m 0,09 % 0,07 % 0,05 % on nominal tank volume | Calibration according to: ISO 7507-3:2006 Measurement - Calculation of Radius: ISO 7507-3:2006, or Measurement of Reference Circumference: ISO 7507-1: 2003 Bottom volume measurement: ISO 4269:2001 or ISO 7507-1:2003 Volumetric table calculation and other measurements: ISO 7507-1:2003 |
|  | Up to 140000 m3 Minimum tank diameter 5 m | Height Diameter 4 m 10 m > 15 m 5 m 0,15 % 0,05 % 0,04 % > 15 m 0,12 % 0,05 % 0,04 % on nominal tank volume | Calibration according to: ISO 7507-4:2010 Bottom volume measurement: ISO 4269:2001 or ISO 7507-1:2003 Volumetric table calculation and other measurements: ISO 7507-1:2003 |
| Volume / Vertical cylindrical tanks (continued) | Up to 140000 m3 Minimum tank diameter 2 m | Height Diameter 3 m 6 m 8 m > 10 m 2 m 0,18 % - - - 4 m - 0,10 % 0,08 % - 10 m - 0,10 % 0,08 % - > 10 m - - - 0,07 % on nominal tank volume | D504-O324 in-house internal electro-optical distance ranging method using 3D laser scanner, based on ISO 7507-4 Volumetric table calculation and other measurements: ISO 7507-1:2003 |
|  | Up to 140000 m3 Minimum tank diameter 1 m | Height Diameter 3 m 6 m 8 m > 10 m 1 m 0,20 % - - - 4 m - 0,10 % 0,08 % - 10 m - 0,10 % 0,08 % - > 10 m - - - 0,07 % on nominal tank volume | D504-O324 in-house external electro-optical distance ranging method using 3D laser scanner, based on ISO 7507-1 Volumetric table calculation and other measurements: ISO 7507-1:2003 |
| Volume / Steel tanks | [5 to 10] m3 | 0,13 % on nominal tank volume | Volumetric method using flow meter according to: ISO 4269: 2001 Calibration medium: potable water. |
|  | (10 to 600] m3 | 0,085 % on nominal tank volume |  |
| Volume / Horizontal cylindrical tanks | Up to 20000 m3 Minimum tank diameter 2 m | Length Diameter 5 m 10 m > 20 m 2 m 0,18 % 0,18 % - 4 m 0,10 % 0,10 % 0,09 % > 12 m - - 0,03 % on nominal tank volume | D504-O320 in-house internal electro-optical distance ranging method using 3D laser scanner, based on: ISO 12917-2:2002 |
|  | Up to 15000 m3 Minimum tank diameter 1 m | Length Diameter 2 m 4 m ≥10 m 1 m 0,44 % 0,43 % 0,42% 2 m 0,24 % 0,22 % 0,21% > 6 m 0,13 % 0,09 % 0,07% on nominal tank volume | D504-O320 in-house external electro-optical distance ranging method using 3D laser scanner, based on: ISO 12917-1:2017 |
| Volume / Spherical tanks | Up to 10000 m3 Minimum tank radius 1,5 m | 0,04 % on nominal tank volume | D504-O318 in-house internal & external electro-optical distance ranging method using 3D laser scanner, based on: API 2552:1966 (R2018) |
| Volume / Ship tanks | [10 to 120] m3 | 0,25 % to 0,12 % on nominal tank volume | D504-O322 in-house electro-optical distance ranging method using 3D laser scanner, based on: API MPMS Chapter 2.8A:1991(R2023) |
|  | (120 to 1500] m3 | 0,12 % to 0,06 % on nominal tank volume |  |
|  | (1500 to 30000] m3 | 0,06 % on nominal tank volume |  |

* 	Where the expanded uncertainty (with 95 % coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration & Measurement Capability (CMC) includes the measured quantity, the measurement range and the measurement uncertainty, expressing the minimum measurement uncertainty which can be achieved in a calibration.

Permanent laboratory premises address: 12th km Athens-Lamia National Road, 14452, Metamorfosi, Attiki, Greece.

Approved Signatories:  D. Gerasimou, A. Vasileiou.

This Scope of Accreditation replaces the previous one dated 18.12.2023.

The Accreditation Certificate No. 364-7, to ELOT EN ISO/IEC 17025:2017, is valid until 9.07.2028.

Athens, 12.02.2024

Christos Nestoras

CEO of ESYD