Hellenic Accreditation System

Annex F2/14 to the Certificate No. 186-7

SCOPE of ACCREDITATION of the Calibration Laboratory Calibration Technologies PC (CPT IKE)

| Calibration item / Parameters | Range of measurement | Expanded Uncertainty of Measurement (k=2)* | Remarks |
| --- | --- | --- | --- |
| Standard weights |  |  | Capability to calibrate standard weights according to OIML R 111: |
|  | 1 mg | 18 μg | up to (including) class F2 |
|  | 2 mg | 18 μg | ,, F2 |
|  | 5 mg | 18 μg | ,, F2 |
|  | 10 mg | 18 μg | ,, F2 |
|  | 20 mg | 18 μg | ,, F2 |
|  | 50 mg | 18 μg | ,, F2 |
|  | 100 mg | 18 μg | ,, F2 |
|  | 200 mg | 19 μg | ,, F1 |
|  | 500 mg | 19 μg | ,, F1 |
|  | 1 g | 19 μg | ,, F1 |
|  | 2 g | 19 μg | ,, F1 |
|  | 5 g | 19 μg | ,, F1 |
|  | 10 g | 19 μg | ,, F1 |
|  | 20 g | 20 μg | ,, F1 |
|  | 50 g | 25 μg | ,, F1 |
|  | 100 g | 40 μg | ,, F1 |
|  | 200 g | 50 μg | ,, F1 |
|  | 500 g | 0,12 mg | ,, F1 |
|  | 1 kg | 1,2 mg | ,, F1 |
|  | 2 kg | 1,3 mg | ,, F1 |
|  | 5 kg | 1,6 mg | ,, F1 |
|  | 10 kg | 2,3 mg | ,, F1 |
|  | 20 kg | 0,3 g | ,, M1 |
|  | 50 kg | 0,9 g | ,, M2 |
| Calibration item / Parameters | Range of measurement | Expanded Uncertainty of Measurement (k=2)* | Remarks |
| Non-automatic electronic weighing instruments | 1 mg | 2,9 μg | Calibration according to guide EIM ΜΜ-ΜΑ-01a v02:2007 and Technical Guide EURAMET cg-18 ver 4.0 Available standards OIML R 111 Class: E2: 1 mg – 5 kg (max grouped load 8,8 kg) F1: 1 g – 50 kg (max grouped load 100 kg) F2: 10 kg (max grouped load 10 kg) M2: 10 kg – 50 kg (max grouped load 400 kg) Calibration is performed on-site |
|  | 2 mg | 2,9 μg |  |
|  | 5 mg | 3,1 μg |  |
|  | 10 mg | 3,1 μg |  |
|  | 20 mg | 3,1 μg |  |
|  | 50 mg | 3,2 μg |  |
|  | 100 mg | 7,3 μg |  |
|  | 200 mg | 7,6 μg |  |
|  | 500 mg | 8,1 μg |  |
|  | 1 g | 9,2 μg |  |
|  | 2 g | 9,0 μg |  |
|  | 5 g | 12 μg |  |
|  | 10 g | 14 μg |  |
|  | 20 g | 26 μg |  |
|  | 50 g | 27 μg |  |
|  | 100 g | 35 μg |  |
|  | 200 g | 0,14 mg |  |
|  | 500 g | 0,20 mg |  |
|  | 1 kg | 0,32 mg |  |
|  | 2 kg | 1,5 mg |  |
|  | 5 kg | 2,0 mg |  |
|  | (5 to 100) kg | 2,4 ∙ 10-6 |  |
|  | (100 to 400) kg | 2,3 ∙ 10-4 |  |
|  | (400 to 4.000) kg | (1,2 kg + 4,50∙10-5∙m)∙Ν | Calibration according to Technical Guide EURAMET cg-18 ver 4.0 Maximum available grouped load of mass standards: 400 kg m: load of mass standards in kg N: ratio of current check load value to the mass standards load (m) Calibration is performed on-site |
| Industrial process weighing systems | (1 to 3000) kg | 1,16∙10-3 | Static calibration according to guide WG00496 of Measurement and Control Institute: “A code of practice for the calibration of industrial process Weighing Systems: 2011” Maximum available grouped load of mass standards: 3000 kg |

| Calibration item / Parameters | Range of measurement | Expanded Uncertainty of Measurement (k=2)* | Remarks |
| --- | --- | --- | --- |
| Compression testing machines | [ 0,5 to 5] kN | 0,13 % | Class 0.5 machines to ΕLΟΤ ΕΝ ISO 7500-1: 2016 |
|  | (5 to 15] kN | 0,22 % | Class 1 machines to ΕLΟΤ ΕΝ ISO 7500-1: 2016 |
|  | (15 to 50] kN | 0,17 % | Class 0.5 machines to ΕLΟΤ ΕΝ ISO 7500-1: 2016 |
|  | (200 to 2000] kN | 0,25 % | Class 1 machines to ΕLΟΤ ΕΝ ISO 7500-1: 2016 |
|  |  |  | Calibration is performed on-site |
| Concrete specimen compression testing machines |  |  | Class 1 machines to ΕΝ 12390-4: 2019 According to ΕΝ 12390-4: 2019 Calibration is performed on-site |
| - Force indication | (200 to 2700) kN | 0.15 % |  |
| - Force transfer | (200 to 2700) kN | 0,01 |  |
| Length/ Calipers: |  |  |
| A) Analog / division 20 μm | (0 to 500) mm | 12 μm | VDI/VDE/DGQ 2618 part 9.1: 2006 or/and ISO 13385-1:2019 |
| B) Analog / division 50 μm | (0 to 500) mm | 29 μm |  |
| C) Analog / division 100 μm | (0 to 500) mm | 57 μm |  |
| D) Digital / division 10 μm | (0 to 500) mm | 6 μm |  |
| Length/ External spindle micrometers: |  |  |  |
| A) Analog & Digital / division 10 μm | (0 to 25) mm | 5,7 μm | VDI/VDE/DGQ 2618 part 10.1: 2001 |
| B) Analog & Digital / division 1 μm | (0 to 25) mm | 0,6 μm |  |
| Length/ Sieves: |  |  |  |
| Wire cloth | (0,020 to 125) mm | (3,1 – 5,8) μm | ASTM E11: 2020 |
| Metal wire cloth | (0,180 to 125) mm | (3,1 – 5,8) μm | ISO 3310-1: 2016 |
| Perforated metal plate (circular holes) | (1 to 125) mm | (3,5 – 5,9) μm | ISO 3310-2: 2013 (except plate thickness and frame size) |
| Perforated metal plate (rectangular holes) | (4 to 125) mm | (3,5 – 5,9) μm |  |

| Calibration item / Parameters | Range of measurement | Expanded Uncertainty of Measurement (k=2)* | Remarks |
| --- | --- | --- | --- |
| Analog & Digital direct reading thermometers | 0 oC | 0,045 oC | Internal calibration method O-50 using liquid bath and temperature resistance probe |
|  | 30 oC to 90 oC | 0,045 oC |  |
| Total immersion, liquid in glass thermometers | 0 oC | 0,045 oC | Internal calibration method O-50 using liquid bath and temperature resistance probe |
|  | 30 oC to 90 oC | 0,045 oC |  |
| Analog and digital room thermometers and temperature data loggers | -20 oC to +80 οC | 0, 42 οC | Internal calibration method O-40 using climatic chamber and temperature resistance probe |
| Temperature chambers of volume up to 2000 l (ovens, incubators, refrigerators, freezers, environmental chambers) | -20 oC to +200 οC | 0,40 oC | According to guide DKD-R 5-7:2018. Calibration can be performed on-site. |
| Analog and digital hygrometers and relative humidity recorders | 40 %rh to 95 %rh at 20 0C | 3,5 % rh | Internal calibration method O-40 using climatic chamber and standard wet – dry bulb hygrometer. |
| Gauge pressure meters, gas medium, direct reading / Analog & Digital | (0.1 to 2) Mpa | 35 hPa | Use of pressure indicating gauge Calibration method based on recommendation DKD R-6-1: 2014 |
|  | (2 to 10) Mpa | 40 hPa |  |
|  | (10 to 40) Mpa | 120 hPa |  |

* Where the expanded uncertainty (with 95 % coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration Measurement Capability (CMC) includes the measured quantity, the measurement range and the measurement uncertainty, expressing the minimum measurement uncertainty which can be achieved in a calibration.

Permanent laboratory premises address: Stavrochori Industrial Area, Kilkis, Greece.

Approved Signatories: E. Galanopoulos, K. Zeimbekis, K. Karambeti, G. Gelastopoulos.

It is noted that the Standard

This Scope of Accreditation replaces the previous one dated 11.04.2023.

The Accreditation Certificate No. 186-7, to ELOT EN ISO/IEC 17025:2017, is valid until 25.11.2024.

Athens, 10.10.2023

Christos Nestoras

CEO of ESYD

# Text boxes

Mass measurements

Force measurements

Dimensional measurements

Temperature measurements

Relative humidity measurements

Pressure measurements