Hellenic Accreditation System S.A.

Annex F2/19 to the Certificate No. 196-7

SCOPE of ACCREDITATION of the Calibration Laboratory of the ‘CQS S.A. – Calibration and Quality Services’

| Measurand / Calibration Ιtem | Range of Μeasurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass measurements |
| Mass / Mass standards | 1 mg | 0,015 mg | Capability to calibrate standard weights up to (including): Class F2 according to OIML |
|  | 2 mg | 0,015 mg | `` F2 `` `` `` |
|  | 5 mg | 0,015 mg | `` F2 `` `` `` |
|  | 10 mg | 0,015 mg | `` F2 `` `` `` |
|  | 20 mg | 0,015 mg | `` F2 `` `` `` |
|  | 50 mg | 0,015 mg | `` F2 `` `` `` |
|  | 100 mg | 0,016 mg | `` F1 `` `` `` |
|  | 200 mg | 0,020 mg | `` F1 `` `` `` |
|  | 500 mg | 0,025 mg | `` F1 `` `` `` |
|  | 1 g | 0,03 mg | `` F1 `` `` `` |
|  | 2 g | 0,04 mg | `` F1 `` `` `` |
|  | 5 g | 0,05 mg | `` F1 `` `` `` |
|  | 10 g | 0,06 mg | `` F1 `` `` `` |
|  | 20 g | 0,08 mg | `` F1 `` `` `` |
|  | 50 g | 0,10 mg | `` F1 `` `` `` |
|  | 100 g | 0,16 mg | `` F1 `` `` `` |
|  | 200 g | 0,3 mg | `` F1 `` `` `` |
|  | 500 g | 2,5 mg | `` F2 `` `` `` |
|  | 1 kg | 1,6 mg | `` F1 `` `` `` |
|  | 2 kg | 30 mg | `` M1 `` `` `` |
|  | 5 kg | 80 mg | `` M1 `` `` `` |
|  | 10 kg | 160 mg | `` M1 `` `` `` |
|  | 20 kg | 300 mg | `` M1 `` `` `` |
|  | 50 kg | 2500 mg | `` M2 `` `` `` |
|  | 100 kg | 5000 mg | `` M2 `` `` `` |
|  | 200 kg | 10 g | `` M2 `` `` `` |

| Measurand / Calibration Ιtem | Range of Μeasurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass / Mass standards | 200 kg | 10 g | `` M2 `` `` `` |
|  | 500 kg | 25 g | `` M2 `` `` `` |
|  | 1000 kg | 50 g | `` M2 `` `` `` |
|  | 2000 kg | 100 g | `` M2 `` `` `` |
| Mass / Non-automatic weighing instruments | 1 mg | 0,0086 mg | Available standards: OIML Class E2: 1 mg – 200 g, max grouped load 610g OIML Class F1: 1 mg – 1 kg, max grouped load 2,1kg OIML Class F2: 1 mg – 5 kg max grouped load 22,2 kg OIML Class M1: 5 kg – 20 kg max grouped load 120 kg OIML Class M1: 20 kg standards max grouped load 1000 kg OIML Class M1: standards 100kg, 200kg, 500kg, 1000kg OIML Class M2: standards 2 x 1000 kg, 10 x 2000 kg max grouped load 22000 kg Calibration is performed on-site, according to EURAMET cg-18 v4.0 (11/2015) “Guidelines on the Calibration of Non-Automatic Weighing Instruments”. All calibrations are performed on-site |
|  | 2 mg | 0,0086 mg |  |
|  | 5 mg | 0,0086 mg |  |
|  | 10 mg | 0,0087 mg |  |
|  | 20 mg | 0,0092 mg |  |
|  | 50 mg | 0,0098 mg |  |
|  | 100 mg | 0,011 mg |  |
|  | 200 mg | 0,012 mg |  |
|  | 500 mg | 0,014 mg |  |
|  | 1 g | 0,016 mg |  |
|  | 2 g | 0,021 mg |  |
|  | 5 g | 0,023 mg |  |
|  | 10 g | 0,028 mg |  |
|  | 20 g | 0,040 mg |  |
|  | 50 g | 0,043 mg |  |
|  | 100 g | 0,07 mg |  |
|  | 200 g | 0,14 mg |  |
|  | 500 g | 0,35 mg |  |
|  | 1 kg | 2,4 mg |  |
|  | [2 to 10] kg | 7·10-6 |  |
|  | 20 kg | 420 mg |  |
|  | [50 to 1.100] kg | 8,6·10-5 |  |
|  | (1.000 to 20.000] kg | d=1 kg 2,7·10-4 |  |
|  | 20.000 kg d=1 kg d=10 kg | 5,4 kg 9,6 kg | According to EURAMET cg-18 v4.0 (11/2015) “Guidelines on the Calibration of Non-Automatic Weighing Instruments”, by use of substitution loads. All calibrations are performed on-site |
|  | 40.000 kg d=1 kg d=10 kg | 10 kg 18 kg |  |
|  | 60.000 kg d=1 kg d=10 kg | 17 kg 25 kg |  |
|  | 80.000 kg d=1 kg d=10 kg | 24 kg 32 kg |  |

| Measurand / Calibration Ιtem | Range of Μeasurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Mass / Calibration of vehicle suspension testers | 0 kg to 1050 kg per wheel | 0,1kg (0,01)% | According to the calibration instructions of the manufacturer: MAHA. The procedure additionally includes tests for scale verification, corner load error and repeatability of the instrument. The calibration can be performed on site. Available standards: OIML Class M1: 1kg - 20 kg max grouped load 1000 kg |
| Force measurements |
| Force / Compression testing machines | [1 to 10] kN | 0,12 % | Class 0,5 machines |
|  | [5 to 50] kN | 0,12 % | Class 0,5 machines |
|  | [10 to 100] kN | 0,12 % | Class 0,5 machines |
|  | [200 to 2000] kN | 0,23 % | Class 1 machines |
|  | [400 to 2000] kN | 0,17 % | Class 0,5 machines |
|  |  |  | According to ELΟΤ ΕΝ ISO 7500-1: 2015. Calibration can be performed on-site. |
| Force / Tension testing machines | [1 to 10] kN | 0,13 % | Class 0,5 machines |
|  | [5 to 50] kN | 0,12 % | Class 0,5 machines |
|  |  |  | According to ELΟΤ ΕΝ ISO 7500-1: 2015. Calibration can be performed on-site. |
| Force / Compression testing machines |  |  | According to EN 12390-4. All calibrations can be performed on-site. |
| - Force transfer | 200 kN | 0,02 |  |
|  | [400 to 2000] kN | 0,01 |  |
| - Rate of application of force | 1 to 100 kN/sec | 1,0 % |  |
| - Flatness of platens | - | 4 μm |  |
| - Hardness of platens | [520 to 800] HLD | 26 HLD |  |
| - Roughness of platens | [0,6 to 4] μm (Ra) | 0,18 μm (Ra) |  |
| Force / Load Rings, CBR, Marshall Loading Machines | [1 to 100] kN | 0,12% | According to ELOT ISO 7500-1: 2015. Calibration can be performed on-site. |
| Torque measurements |
| Toque tools, Torque wrenches (setting, fixed torque, indication) | 6 to 60 Nm 20 to 200 Nm 50 to 1000 Nm | 0,2 % 0,3 % 0,4 % | According to ISO 6789: 2003, with standard torque meter |
| Hardness measurements |
| Hardness meters Leeb D and Rockwell C | 520 to 800 HLD 24 to 65 HRC | 7 to 10 HLD 1,3 to 0,45 HRC | Calibrations can be performed on-site. |

| Measurand / Calibration Ιtem | Range of Μeasurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Pressure measurements |
| Pressure / Analog and digital gauge pressure measuring instruments of direct reading | (0 to 4) MPa | (0,48 to 1,4) kPa | Gas pressure medium, according to guide DKD-R 6-1 03/2014. The calibration can be performed on site. |
|  | (0 to 60) MPa | 0,024 MPa | Oil or water pressure medium, according to DKD-R 6-1 03/2014. The calibration can be performed on site. |
|  | (-0,1 to 0,15) MPa | 0,10 kPa | Gas pressure medium, according to guide DKD-R 6-1 03/2014. |
| Dimensional Measurements |
| Length / Digital&analog Calipers | (0 to 1000) mm / d = 10 μm d = 20 μm d = 50 μm d = 100 μm | 6 μm 12 μm 29 μm 58 μm | According to guide VDI/VDE/DGΟ 2618, part 9.1: 2006 |
| Length / Digital & analog external axis Micrometers | (0 to 25) mm / d = 1 μm d = 10 μm | 0,6 μm 6 μm | According to guide VDI/VDE/DGΟ 2618, part 10.1:2001 Using standard gauge blocks and horizontal measuring instrument TRIMOS |
|  | (25 to 125) mm / d = 1 μm d = 10 μm | 0,7 μm 6 μm |  |
|  | (125 to 500) mm / d = 10 μm d = 100 μm | 7 μm 58 μm |  |
| Length/ Horizontal length measuring instrument (TRIMOS) | (0 to 500) mm/ d= 1μm | 1μm - 3μm | Based on guide VDI/VDE/DGΟ 2618, part 9.1:2006 &10.1: 2001 |
| Dial gauge and digital indicators for linear measurement, displacement transducers | 0mm to 100mm / d = 1 μm d = 10 μm d = 100 μm | 3,5 μm 7 μm 58 μm | Comparative measurement Calibration at laboratory |
|  | 0mm to 100mm / d = 1 μm d = 10 μm d = 100 μm | 1 μm 6 μm 58 μm | Using standard gauge blocks. The calibration can be performed on site. |
| Moulds for concrete specimens (Cubic, cylindrical, prismatic) | Internal dimensions Perpendicularity Flatness | 9 μm 7 μm 4 μm | According to EN 12390-1:2012 Check pass / fail Check pass / fail |
| Cylinder moulds (Proctor) | Internal diameter Height | 9 μm 9 μm | According to EN 13286-2:2010 |
| Cylinder moulds (CBR) | Internal diameter Height | 9 μm 9 μm | According to EN 13286-47:2012 |
| Cylinder moulds (Marshall) | Internal diameter | 9 μm | According to EN 12697-30:2012 |
| Measuring Rulers, Tape meters, Level measuring rods | [3·(n-1) to 3·n] m n = 1, 2, 3, … | 50(n+1,4)μm 0,6d (mm) (d ≥ 2mm) | OIML R35-2 §7.5, §7.6 |
| Magnetostrictive level gauges | (0 to 3,5) m / d = 0,1 mm d = 1 mm | 0,13 mm 0,6 mm | Based on OIML R35-2 §7.5, §7.6 |
| Measuring projectors / Scale microscopes | 2 to 300mm d=0,0001mm d-0.001mm | 1,7 μm 7,5 μm (x10 < lens < x100) | Internal method using standard scales. The calibration can be performed on site. |
|  | 0 to 2mm d=0,0001mm d-0.001mm | 0,95 μm 7,5 μm (x10 < lens < x100) |  |

| Measurand / Calibration Ιtem | Range of Μeasurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Sieves | 45μm to 2000mm 1mm to 125mm | 1,3μm (with lens x100) 11μm (with lens x10) | According to standards; ISO 3310-1:2016 ISO 3310-2:2013 ASTM E11 – 17:2020 EN 933-3:2012 ISO 5223:2012 |
| Frequency / Time Measurements |
| Rotation frequency / Reciprocating movement frequency Centrifuges, Extractors, Mixers, Shakers, Vibrating tables | [50 to 1000) rpm | 0,1 rpm | Internal method using standard tachometer. Calibration can also be performed on-site f: rotation frequency in rpm |
|  | [1.000 to 5.000) rpm | 0,6 rpm |  |
|  | [5.000 to 60.000] rpm | 1,8 rpm + 0,01% x f |  |
| Digital and analog stopwatches / timers | Up to 10hr | 0,32sec | Comparative calibration using standard chronometer. The calibration can be performed on site. |
| Temperature Measurements |
| Temperature / Analog, digital thermometers of direct reading, data loggers | [-30 to –0) °C | 0,062 °C | Comparative calibration using liquid bath and temperature resistance probes Calibration can also be performed on-site. |
|  | 0 °C | 0,040 °C |  |
|  | (0 to 150] °C | 0,046 °C |  |
|  | (150 to 250] °C | 0,050 °C |  |
|  | (250 to 400] °C | 0,26 °C | Comparative calibration using using dry block calibrator and temperature resistance probe |
|  | (400 to 500] °C | 0,30 °C |  |
|  | (500 to 600] °C | 0,35 °C |  |
| Temperature / Total and Partial immersion liquid-in-glass thermometers | [-30 –0) °C | 0,062 °C | Comparative calibration using liquid bath and temperature resistance probes. |
|  | 0 °C | 0,040 °C |  |
|  | (0 to 150] °C | 0,046 °C |  |
|  | (150 to 250] °C | 0,050 °C |  |
| Temperature / Thermometers of digital depth meters | [5 to 50] °C | 0,08 °C | Comparative calibration using isothermal configuration and Pt100 probes. |
| Temperature / Analog, Digital thermometers of direct reading, Data loggers | [-40 to 0) °C | 0,17 °C | Comparative calibration using climatic chamber and temperature resistance probes Pt100 |
|  | (0 to 180] °C | 0,29 °C |  |
| Temperature / Platinum Resistance Thermometers Sensors (iPRT) | [-30 to 0) °C | 0,077 °C | Measurement of probe resistance. Using temperature resistance probes Pt100. According to guide DKD-R 5-1 02/2009 |
|  | 0 °C | 0,032 °C |  |
|  | (0 to 150] °C | 0,037 °C |  |
|  | (150 to 250] °C | 0,042 °C |  |
|  | (250 to 400] °C | 0,26 °C | Comparative calibration using dry block calibrator and standard thermometer |

| Measurand / Calibration Ιtem | Range of Μeasurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Temperature / Climatic and isothermal chambers,Ovens, Incubators, Refrigerators, Freezers, Autoclaves, Isothermal liquid baths with volume up to 2000 l | [-20 to .150) oC | 0,19oC | According to guide DKD-R 5-7 07/2004, and ELΟΤΕΝ 60068-3-5:2002, using temperature resistance probes Pt 100. Calibration is carried out in the useful volume or in selected locations in the chamber. Calibration can be performed on-site. |
|  | (150 to 250] oC | 0,6 oC |  |
| Temperature / Climatic and isothermal chambers,Ovens, Incubators, Refrigerators, Freezers, Autoclaves, High Temperature Furnaces, Isothermal liquid baths with volume up to 2000 l | [-20 to .150) oC | 0,35 oC | According to guide DKD-R 5-7 07/2004, and ELΟΤΕΝ 60068-3-5:2002, using type K thermocouples. Calibration is carried out in the useful volume or in selected locations in the chamber. Calibration can be performed on-site. |
|  | (150 to 250] oC | 0,8 oC |  |
|  | (250 to 500] oC | 1,2oC |  |
|  | (500 to 600] oC | 1,3oC | Using type K thermocouples. Calibration is carried out in the useful volume or in selected locations in the chamber. Calibration can be performed on-site. |
|  | (600 to 1200] °C | 1,3 °C | Using type K thermocouples. Calibration is carried out in selected locations in the chamber. Calibration can be performed on-site. |
| Temperature / Climatic and isothermal chambers, Ovens, Incubators, Refrigerators, Freezers, with volume greater than 2000 l | [-20 to 60] °C | 0,35 °C | According to ELOT ΕΝ 60068-3-5:2002, using digital temperature recorders. Calibration is carried out in the useful volume or in selected locations in the chamber. Calibration can be performed on-site. |
| Temperature / Thermal cyclers (PCR), Apparatus for determining COD, Kjeldahl devices, Heating blocks | [-40 to +150) oC | 0,3 oC | According to guide EURAMET/cg-13/v.3.0 (02/2015), using temperature resistance probes Pt100. The calibration is carried out in the entire block or at selected points. Calibration can also be performed on-site. |
|  | (+150 to +250] oC | 0,4 oC |  |
| Temperature / Thermal cyclers (PCR), Apparatus for determining COD, Kjeldahl devices, Heating blocks | [-40 to +100) oC | 0,3 oC | According to guide EURAMET/cg-13/v.3.0 (02/2015), using type K thermocouples. The calibration is carried out in the entire block or at selected points. Calibration can also be performed on-site. |
|  | (+100 to +250) oC | 0,4 oC |  |
|  | (+250 to +420] oC | 0,9 oC |  |
|  | (+420 to 600] οC | 1,2 oC |  |
| Temperature / Dissolution testers and disintegration testers | (+30 to +50) °C | 0,3 oC | Using temperature resistance probes Pt100 |
| Temperature / Dryblock calibrators | [-80 to +250) °C | 0,2 oC | Using temperature resistance probes Pt100 according to guide EURAMET/cg-13/v.3.0 (02/2015) |
|  | (+250 to + 400] °C | 0,32 oC |  |
| Measurand / Calibration Ιtem | Range of Μeasurement | Calibration & Measurement Capability (k=2)* | Remarks |
| Relative Humidity |
| Relative humidity chambers with volume up to 2000 l | (30 to 98) % RH at 23οC | 1,2 %RH | According to guide DKD-R 5-7 (2004) |
| Relative humidity / Analog and digital instruments measuring and logging relative humidity of direct reading. | [30 to 70) % RH at 23οC | 1,5 %RH | Comparative calibration using climatic chamber and standard hygrometer. |
|  | (70 to 98] % RH at 23οC | 2,0 %RH |  |
| Spectrophotometric measurements |
| Absorption (Α) / UVVis spectrophotometers in the range 235-635nm | visible spectrum (Α ≤ 0,5 AU) | 2,8 mAU ** | ASTME 275:2008 (Re 2013) Calibration can be performed on-site. |
|  | visible spectrum (0,5 < Α ≤ 1,0 AU) | 3,3 mAU ** |  |
|  | UV spectrum (Α ≤ 0,5 AU) | 3,8 mAU ** |  |
|  | UV spectrum (0,5 <Α ≤ 1,5 AU) | 6,0 mAU ** |  |
| Volume Measurements |
| Volume / Piston operated measuring pipettes | 2μl | 5,00 % | Following the standard EURAMET cg-19 v 3.0 (09/2018) “Guidelines on the determination of uncertainty in gravimetric volume calibration” |
|  | 10 μl | 1,10 % |  |
|  | 20 μl | 0,55 % |  |
|  | 100 μl | 0,18 % |  |
|  | 1000 μl | 0,14 % |  |
|  | 5 ml | 0,12 % |  |
|  | 10 ml | 0,12 % |  |
| Volume / Piston operated volume measuring instruments –Dispensers, Dilutors, Piston burettes | 0,5 ml | 0,12 % |  |
|  | 1 ml | 0,10 % |  |
|  | 2 ml | 0,10 % |  |
|  | 5 ml | 0,08 % |  |
|  | 10 ml | 0,05 % |  |
|  | 25 ml | 0,05 % |  |
|  | 50 ml | 0,05 % |  |
|  | 100 ml | 0,04 % |  |
|  | 200 ml | 0,04 % |  |
| Volume / Burettes | [1 to 100]ml | [0,20 to 0,04] % |  |
| Volume / Volumetric flasks | 5 ml | 0,16 % |  |
|  | 10 ml | 0,09 % |  |
|  | 100 ml | 0,03 % |  |
|  | 1000 ml | 0,02 % |  |
|  | 2000 ml | 0,02 % |  |
| Volume / Volumetric Cylinders | 5 ml | 0,22 % |  |
|  | 10 ml | 0,18 % |  |
|  | 50 ml | 0,12 % |  |
|  | 100 ml | 0,08 % |  |
|  | 250 ml | 0,08 % |  |
|  | 500 ml | 0,08 % |  |
|  | 1000 ml | 0,05 % |  |
|  | 2000 ml | 0,05 % |  |
| Volume / Bulb pipettes | [0,5 to 100] ml | [0,16 to 0,03] % |  |

| Measurand / Calibration Ιtem | Range of Μeasurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Volume/ Volumetric vessels «IN» of “slicker plate” and “neck scale” type | 2 l | 0,11 % | EURAMET cg-19 v3.0 (09/2018) “Guidelines on the determination of uncertainty in gravimetric volume calibration” |
|  | 20 l | 0,04 % |  |
|  | 100 l | 0,04 % |  |
| Liquid Volume / Stationary metal tanks (horizontal) | 0,5 to 100 m3 | 0,11% of nominal tank volume | Method followed: using reference flow meter Calibration medium: Tap water Standards followed: ΕΙΜ: Calibration Guidance / Tank volume measurement with volumetric method EURAMET cg-21v1.0 (04/2013) “Guidelines on the determination of uncertainty in gravimetric volume calibration” The calibration can be performed on site. |
| Liquid Volume / Stationary metal tanks (vertical) | 0,2 to 100 m3 | 0,11% of nominal tank volume | Volumetric method using volumetric proving tank or reference flow meter. Calibration medium: Tap water EURAMET cg-21 v1.0 (04/2013) “Guidelines on the determination of uncertainty in gravimetric volume calibration” FEK (Leaf of the Government Gazεττε) Issue Β’ 914 / 15.03.2018 The calibration can be performed on site. |

*Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

**Uncertainty values were estimated assuming a spectrophotometer having a resolution of 0,1 mAU.

Site of assessment: Permanent laboratory premises, 44 Αg.Triados Str. - Pefka, Thessaloniki, Greece.

Approved Signatories: A. Kougioumtzidis, C. Vratmiras, S. Fotiadis, V. Stravas, A. Kyriazos.

This Scope of Accreditation replaces the previous one dated 31.05.2021.

The Accreditation Certificate No. 196-7, to ELOT EN ISO/IEC 17025:2017, is valid until 22.03.2025.

Athens, May 4th 2023

Christos Nestoras

CEO of ESYD