Hellenic Accreditation System

Annex F2/19  to the Certificate No. 383-5

SCOPE of ACCREDITATION of the Calibration Laboratory of Analytical Instruments S.Α.

| Calibration item/ Parameters | Range of measurement | Expanded measurement uncertainty (k=2)* | Remarks |
| --- | --- | --- | --- |
| Spectrophotometric measurements |
| Visible and ultraviolet spectrophotometers | Visible spectrum 0 to 1.5 AU | 8.1 mAU | According to guides: ASTM E 275-08(2022) European Pharmacopoeia method 2.2.25 Absorption Spectrophotometry, Ultraviolet and Visible (Edition 10) Range of measurement: 230 nm to 400nm (ultraviolet) and 400nm to 700 nm (visible). Calibration is also performed on-site |
|  | UV spectrum 0 to 1.5 AU | 5.8 mAU |  |
| Temperature measurements |
| Analog & digital thermometers of direct reading | -20οC to 200οC | 0,16οC | According to COFRAC guide: LAB GTA 08:2017 Using liquid bath and platinum resistance thermometer |
|  | 0οC | 0,09οC | According to COFRAC guide: LAB GTA 08:2017 Using ice-water bath and platinum resistance thermometer |
|  | 200οC to 350οC | 0,92οC | According to COFRAC guide: LAB GTA 08: 2017 Using dry block calibrator and platinum resistance thermometer |
| Climatic and isothermal chambers, ovens, refrigerators, furnaces, incubators, freezers / Internal volume up to 2000 l | -20οC to 100οC | 0,47οC | According to guides: EN 60068-3-5:2018 DKD-R 5-7:2018 using thermocouples type K (methods A,B,C) Range 200οC to 350οC is covered only by method C Calibration is also performed on-site |
|  | 100οC to 200οC | 0,70οC |  |
|  | 200οC to 350οC | 1,2οC |  |
| Relative humidity measurements |
| Relative humidity chambers with air circulation systems of internal volume up to 2000 l (climatic chambers) | [35 to 95) % RH at 25οC | 1,7 % RH | According to guide DKD-R 5-7:2018 (methods A,B,C) |
| Mass measurements |
| Non-automatic weighing instruments | 10 g | 3,5(10-6 | According to guide EURAMET cg18 (v4.0) 2015 Available standard weights E2 OIML class: 1 mg to 200 g, 50 g, 100 g, 2 x 500 g, 3 x 1000 g, 2000 g Calibration is performed on site |
|  | 20 g | 2,3(10-6 |  |
|  | 50 g | 1,1(10-6 |  |
|  | 100 g | 9,0(10-7 |  |
|  | 200 g | 8,9(10-7 |  |
|  | 500 g | 9,3(10-7 |  |
|  | 1000 g | 9,0(10-7 |  |
|  | 2000 g | 9,0(10-7 |  |
|  | 5000 g | 9,1(10-7 |  |
|  | 5500 g | 9,0(10-7 |  |
| Volume measurements |
| Piston operated pipettes | 20 μl to 100 μl | 0,50 % | According to guides: ISO 8655-6: 2022 ISO/TR 20461:2000 |
|  | 100 μl to 1000 μl | 0,25 % |  |
|  | 1 ml to 10 ml | 0,15 % |  |

* Where the expanded uncertainty (with 95 % coverage) is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

The Calibration Measurement Capability (CMC) includes the measured quantity, the measurement range and the measurement uncertainty, expressing the minimum measurement uncertainty which can be achieved in a calibration.

Permanent laboratory premises address: 9, Tzavella Str., 152 31 Chalandri.

Approved Signatories:  Charalampos Perdikis.

This Scope of Accreditation replaces the previous one dated 06.09.2023.

The Accreditation Certificate No. 383-5, to ELOT EN ISO/IEC 17025:2017, has been extended according to ESYD’s decision and it is valid until 23.06.2024.

Athens, 21.03.2024

Christos Nestoras

CEO of ESYD