Hellenic Accreditation System

Annex F2/16 to the Certificate No. 39-9

SCOPE of ACCREDITATION of the Calibration Laboratory of INTRACOM DEFENSE S.A.

| Measurand / calibration item | Range of measurement | Calibration & Measurement Capability (k=2)* | Remarks |
| --- | --- | --- | --- |
| Electrical measurements |
| DC Voltage (Generate – Measure) / DC Voltage Standards, DC Voltage Calibrators & Measuring Equipment | [1 μV to 10 μV] | 0,05 μV | Internal method DCV-A2 |
|  | (10 μV to 1mV] | 0,1 mV |  |
|  | (1 mV to 120 mV) | 1,0 mV |  |
|  | [120 mV to 1 V) | 3,5x10-6 |  |
|  | 1,018 V | 8,0x10-7 |  |
|  | 10 V | 5,0x10-7 |  |
|  | [1 V to 1000 V] | 2,0x10-6 |  |
| DC Current (Generate – Measure) / DC Current Calibrators & Measuring Equipment | [10 μA to 5 mA] | 1,2x10-5 | Internal method DCI-A1 |
|  | (5 mA to 200 mA] | 3,0x10-5 |  |
|  | (200 mA to 3 A] | 3,5x10-5 |  |
|  | (3 A to 20 A] | 7,5x10-5 |  |
| DC Current (Measure) | (20 Α to 100 Α] | 7,5x10-5 |  |
| DC Resistance / Calibration of resistors | [0,5 mW to 1 mW) | 9,5x10-5 | Internal method DCR-A1 |
|  | [1 mW to 3 mW) | 5,5x10-5 |  |
|  | [3 mW to 10 mW) | 3,0x10-5 |  |
|  | [10 mW to 100 mW) | 2,0x10-5 |  |
|  | [100 mW to 10 W] | 9,0x10-6 |  |
|  | (10 W to 100 W) | 1,2x10-5 |  |
|  | [100 W to 1 kW) | 9,0x10-6 |  |
|  | [1 kW to 100 kW] | 7,0x10-6 |  |
|  | (100 kW to 1 MW] | 1,6x10-5 |  |
|  | (1 MW to 10 MW] | 6,0x10-5 |  |
|  | [0,5 mW to 1 mW) | 9,5x10-5 |  |
| DC resistance / Calibration of resistance meters | 10 mW | 2,0x10-5 |  |
|  | 33 mW | 2,2x10-5 |  |
|  | 100 mW | 2,0x10-5 |  |
|  | 0,2 W | 1,5x10-5 |  |
|  | 1 W | 7,0x10-6 |  |
|  | 1,9 W | 1,0x10-4 |  |
|  | 10 W | 1,5x10-5 |  |
|  | 19 W | 2,8x10-5 |  |
|  | 100 W, 1 kW | 1,0x10-5 |  |
|  | 190 W | 1,8x10-5 |  |
|  | 1,9 kW | 1,3x10-5 |  |
|  | 10 kW | 8,0x10-6 |  |
|  | 19 kW | 1,2x10-5 |  |
|  | 100 kW | 1,0x10-5 |  |
|  | 190 kW | 1,4x10-5 |  |
|  | 1 MW | 1,6x10-5 |  |
|  | 1,9 MW | 2,1x10-5 |  |
|  | 10 MW | 5,0x10-5 |  |
|  | 19 MW | 5,0x10-5 |  |
|  | 100 MW | 1,1x10-4 |  |
| AC Voltage (Generate – Measure) / AC Voltage Calibrators & Measuring Equipment AC Voltage (Generate – Measure) / AC Voltage Calibrators & Measuring Equipment (cont.) AC Voltage (Generate – Measure) / AC Voltage Calibrators & Measuring Equipment (cont.) | 100 mV | 10 Hz | 0,16x10-3 | Internal method ACV-A1 |
|  |  | 20 Hz, 40 Hz | 60x10-6 |  |
|  |  | 500 Hz, 1 kHz | 50x10-6 |  |
|  |  | 10 kHz, 20 kHz | 5x10-6 |  |
|  |  | 50 kHz, 70 kHz, 100 kHz | 70x10-6 |  |
|  |  | 200 kHz, 500 kHz | 0,16x10-3 |  |
|  |  | 700 kHz | 0,45x10-3 |  |
|  |  | 1 MHz | 0,72x10-3 |  |
|  | 200 mV | 10 Hz | 0,16x10-3 |  |
|  |  | 20 Hz, 40 Hz | 45x10-6 |  |
|  |  | 500 Hz, 1 kHz | 40x10-6 |  |
|  |  | 10 kHz, 20 kHz | 40x10-6 |  |
|  |  | 50 kHz, 70 kHz, 100 kHz | 50x10-6 |  |
|  |  | 200 kHz | 75x10-6 |  |
|  |  | 500 kHz | 95x10-6 |  |
|  |  | 700 kHz | 0,39x10-3 |  |
|  |  | 1 MHz | 0,66x10-3 |  |
|  | 600 mV | 10 Hz | 65x10-6 |  |
|  |  | 20 Hz, 40 Hz | 45x10-6 |  |
|  |  | 500 Hz, 1 kHz | 35x10-6 |  |
|  |  | 10 kHz, 20 kHz, 50 kHz | 35x10-6 |  |
|  |  | 70 kHz, 100 kHz | 35x10-6 |  |
|  |  | 200 kHz | 45x10-6 |  |
|  |  | 500 kHz | 75x10-6 |  |
|  |  | 700 kHz | 95x10-6 |  |
|  |  | 1 MHz | 0,25x10-3 |  |
|  | 1 V | 10 Hz | 65x10-6 |  |
|  |  | 20 Hz | 35x10-6 |  |
|  |  | 40 Hz, 500 Hz, 1 kHz | 25x10-6 |  |
|  |  | 10 kHz, 20 kHz, 50 kHz | 25x10-6 |  |
|  |  | 70 kHz | 30x10-6 |  |
|  |  | 100 kHz | 35x10-6 |  |
|  |  | 200 kHz | 45x10-6 |  |
|  |  | 500 kHz | 75x10-6 |  |
|  |  | 700 kHz | 95x10-6 |  |
|  |  | 1 MHz | 0,25x10-3 |  |
|  | 2 V | 10 Hz | 65x10-6 |  |
|  |  | 20 Hz | 35x10-6 |  |
|  |  | 40 Hz | 25x10-6 |  |
|  |  | 500 Hz, 1 kHz | 20x10-6 |  |
|  |  | 10 kHz, 20 kHz | 20x10-6 |  |
|  |  | 50 kHz | 25x10-6 |  |
|  |  | 70 kHz | 30x10-6 |  |
|  |  | 100 kHz | 35x10-6 |  |
|  |  | 200 kHz | 45x10-6 |  |
|  |  | 500 kHz | 75x10-6 |  |
|  |  | 700 kHz | 95x10-6 |  |
|  |  | 1 MHz | 0,25x10-3 |  |
|  | 6 V | 10 Hz | 65x10-6 |  |
|  |  | 20 Hz, | 35x10-6 |  |
|  |  | 40 Hz, | 25x10-6 |  |
|  |  | 500 Hz, 1 kHz | 20x10-6 |  |
|  |  | 10 kHz, 20 kHz, 50 kHz | 20x10-6 |  |
|  |  | 70 kHz, 100 kHz | 25x10-6 |  |
|  |  | 200 kHz | 35x10-6 |  |
|  |  | 500 kHz | 45x10-6 |  |
|  |  | 700 kHz | 75x10-6 |  |
|  |  | 1 MHz | 0,25x10-3 |  |
|  | 10 V, 20 V | 10 Hz | 65x10-6 |  |
|  |  | 20 Hz, | 35x10-6 |  |
|  |  | 40 Hz, | 30x10-6 |  |
|  |  | 500 Hz, 1 kHz | 25x10-6 |  |
|  |  | 10 kHz, 20 kHz, 50 kHz | 25x10-6 |  |
|  |  | 70 kHz, 100 kHz | 25x10-6 |  |
|  |  | 200 kHz | 45x10-6 |  |
|  |  | 500 kHz | 75x10-6 |  |
|  |  | 700 kHz | 95x10-6 |  |
|  |  | 1 MHz | 0,25x10-3 |  |
|  | 60 V | 10 Hz | 65x10-6 |  |
|  |  | 20 Hz, 40 Hz | 35x10-6 |  |
|  |  | 500 Hz, 1 kHz | 25x10-6 |  |
|  |  | 10 kHz, 20 kHz, 50 kHz | 25x10-6 |  |
|  |  | 70 kHz | 30x10-6 |  |
|  |  | 100 kHz | 35x10-6 |  |
|  | 100 V, 200 V | 10 Hz | 65x10-6 |  |
|  |  | 20 Hz, 40 Hz | 35x10-6 |  |
|  |  | 500 Hz, 1 kHz, 10 kHz | 25x10-6 |  |
|  |  | 20 kHz, | 30x10-6 |  |
|  |  | 50 kHz 70 kHz | 35x10-6 |  |
|  |  | 100 kHz | 45x10-6 |  |
|  | 1000 V | 10 Hz | 65x10-6 |  |
|  |  | 20 Hz, 40 Hz, 500 Hz | 55x10-6 |  |
|  |  | 1 kHz, 10 kHz, 20 kHz | 55x10-6 |  |
|  |  | 50 kHz | 55x10-6 |  |
|  |  | 70 kHz | 55x10-6 |  |
|  |  | 100 kHz | 0,12x10-3 |  |
| AC Current (Generate – Measure) / AC Current Calibrators & Measuring Equipment | [100 mA to 200 mA] | 40 Hz, 1 kHz, 5 kHz, 10 kHz | 1,5x10-4 | Internal method ACI-A1 |
|  | (200 mA to 300 mA] |  | 3,0x10-4 |  |
|  | (300 mA to 500 mA] |  | 1,9x10-4 |  |
|  | (0,5 A to 1 A] |  | 1,6x10-4 |  |
|  | (1 A to 2 A] |  | 1,9x10-4 |  |
|  | (2 A to 10 A] |  | 3,0x10-4 |  |
|  | (10 A to 20 A] | 40 Hz, 1 kHz, 5 kHz | 3,0x10-4 |  |
| Frequency measurements |
| Frequency f / Measuring instruments & sources of frequency Frequency f / Measuring instruments & sources of frequency (cont) | [1 Hz to 3 GHz] | Digital frequency measurement | uTf: Trigger uncertainty | Internal method FREQ-A1 |
|  | (3 GHz to 26,5 GHz] |  | 1 Hz |  |
|  | 1MHz to 10 MHz | Measurement of phase time difference, 24 h average, range in 1 MHz steps. | 8×10-12 f |  |
| Time interval measurements |
| Time / Time Interval t | 5 ns to 10000 s | uTt: Trigger uncertainty | Internal method FREQ-A1 |
| Temperature measurements |
| Temperature / Thermometers (direct reading) | -40 (C to 120 (C | 45 mK | Internal method Therm_Meas_A1 using temperature calibrator and platinum resistor thermometer |
| Torque measurements |
| Torque / Torque wrenches, Torque drivers | [0,2 Nm to 2 Nm) clockwise only (CW) | 1,3 % | According to standards: ISO 6789-1:2017 ISO 6789-2:2017 using torque measurement system |
|  | [2 Nm to 20 Nm] | 0,60 % |  |
|  | (20 Nm to 100 Nm] | 0,65 % |  |
|  | (100 Nm to 1000 Nm] | 0,70 % |  |

* Where uncertainty is accompanied by the corresponding unit, it is absolute, while where it is not accompanied by a unit, it is relative.

Site of assessment: Building C2, 21 km Markopoulou Ave, 19441 Koropi, Attiki, Greece.

Approved Signatories: E. Vlachakis, G. Lagonikos.

This Scope of Accreditation replaces the previous one dated 24.12.2021.

The Accreditation Certificate No. 39-9, to ELOT EN ISO/IEC 17025: 2017, is valid until 11.11.2025.

Athens, 23.05.2023

Christos Nestoras

CEO of ESYD